Correlation between:

ISO 13485:2003

US Quality System Regulation

Japan GMP

| ISO 13485:2003 | U.S. Quality System Regulation (21 CFR 820) | Japan GMP | Comments |
| --- | --- | --- | --- |
| 1 Scope 1.1 General This International Standard specifies requirements for a quality management system where an organization needs to demonstrate its ability to provide medical devices and related services that consistently meet customer requirements and regulatory requirements applicable to medical devices and related services. The primary objective of this International Standard is to facilitate harmonized medical device regulatory requirements for quality management systems. As a result, it includes some particular requirements for medical devices and excludes some of the requirements of ISO 9001 that are not appropriate as regulatory requirements. Because of these exclusions, organizations whose quality management systems conform to this International Standard cannot claim conformity to ISO 9001 unless their quality management systems conform to all the requirements of ISO 9001 (see Annex B) | 820.1 Scope Applicability. (1) Current good manufacturing practice (CGMP) requirements are set forth in this quality system regulation. The requirements in this part govern the methods used in, and the facilities and controls used for, the design, manufacture, packaging, labeling, storage, installation, and servicing of all finished devices intended for human use. The requirements in this part are intended to ensure that finished devices will be safe and effective and otherwise in compliance with the Federal Food, Drug, and Cosmetic Act (the act). This part establishes basic requirements applicable to manufacturers of finished medical devices. … (c) Authority. Part 820 is established and issued under authority of sections 501, 502, 510, 513, 514, 515, 518, 519, 520, 522, 701, 704, 801, 803 of the act (21 U.S.C. 351, 352, 360, 360c, 360d, 360e, 360h, 360i, 360j, 360l, 371, 374, 381, 383). The failure to comply with any applicable provision in this part renders a device adulterated under section 501(h) of the act. Such a device, as well as any person responsible for the failure to comply, is subject to regulatory action. (d) Foreign manufacturers. If a manufacturer who offers devices for import into the United States refuses to permit or allow the completion of a Food and Drug Administration (FDA) inspection of the foreign facility for the purpose of determining compliance with this part, it shall appear for purposes of section 801(a) of the act, that the methods used in, and the facilities and controls used for, the design, manufacture, packaging, labeling, storage, installation, or servicing of any devices produced at such facility that are offered for import into the United States do not conform to the requirements of section 520(f) of the act and this part and that the devices manufactured at that facility are adulterated under section 501(h) of the act. (e) Exemptions or variances. (1) Any person who wishes to petition for an exemption or variance from any device quality system requirement is subject to the requirements of section 520(f)(2) of the act. Petitions for an exemption or variance shall be submitted according to the procedures set forth in Sec. 10.30 of this chapter, the FDA's administrative procedures. Guidance is available from the Center for Devices and Radiological Health, Division of Small Manufacturers Assistance, (HFZ-220), 1350 Piccard Dr., Rockville, MD 20850, U.S.A., telephone 1-800-638-2041 or 1-301-443-6597, FAX 301-443-8818. (2) FDA may initiate and grant a variance from any device quality system requirement when the agency determines that such variance is in the best interest of the public health. Such variance will remain in effect only so long as there remains a public health need for the device and the device would not likely be made sufficiently available without the variance. | Scope Article 3 The marketing authorisation holder of the medical devices or in-vitro diagnostic reagents specified in Paragraph 1 of Article 14 of Law, the appointed marketing authorisation holder of medical devices or in-vitro diagnostic reagents specified in Paragraph 4 of Article 19-2 of Law or the marketing authorisation holder of the designated controlled medical devices specified, etc. in Paragraph 1 of Article 23-2 of Law (hereinafter collectively referred to as “marketing authorisation holder, etc.”) shall, in accordance with the provision of Chapter 2 or the provision of Chapter 2 applied mutatis mutandis under Chapter 5, have the manufacturer, etc. conduct the manufacturing control and quality control of the products in the manufacturing site, with the proviso that the manufacturing control and quality control in the manufacturing site of the manufacturer in the category specified in Item (4) of Paragraph 5 of Article 26 of Pharmaceutical Affairs Law Enforcement Regulations (MHW Ministerial Ordinance No. 1, 1961) (hereinafter referred to as “Enforcement Regulations”) (hereinafter referred to as “labelling, etc.-category medical device manufacturer”) or the foreign manufacturer in the category specified in Item (4) of Paragraph 4 of Article 36 of Enforcement Regulations (hereinafter referred to as “foreign labelling, etc.-category medical device manufacturer”) (hereinafter collectively referred to as “labelling, etc.-category medical device manufacturer, etc.”), or the manufacturer in the category specified in Item (3) of Paragraph 2 of Article 26 of Enforcement Regulations (hereinafter referred to as “labelling, etc.-category in-vitro diagnostic reagents manufacturer”) or the foreign manufacturer in the category specified in Item (3) of Paragraph 2 of Article 36 of Enforcement Regulations (hereinafter referred to as “foreign labelling, etc.-category in-vitro diagnostic reagents manufacturer”) (hereinafter collectively referred to as “labelling, etc.-category in-vitro diagnostic reagents manufacturer, etc.”) are conducted in accordance with the provision of Chapter 3 or the provision of Chapter 3 applied mutatis mutandis under Chapter 5 may be applied, instead of the provision of Chapter 2 or the provision of Chapter 2 applied mutatis mutandis under Chapter 5. The marketing authorisation holder, etc. of medical devices shall have the manufacturer of the products concerned with the medical devices which correspond to the biological-origin products specified in Paragraph 9 of Article 2 of Law (hereinafter referred to as “biological-origin medical devices”), the medical devices designated by Minister of Health, Labour and Welfare in accordance with the provision of Paragraph 2 of Article 43 of Law, or the cell/tissue-based medical devices (hereinafter collectively referred to as “biological-origin medical devices, etc.”) (hereinafter referred to as 5 “biological-origin medical devices, etc. manufacturer”), and the foreign manufacturer of the products concerned with the biological-origin medical devices, etc. (hereinafter collectively referred to as “biological-origin medical devices, etc. manufacturer, etc.”) conduct the manufacturing control and quality control of the products in the manufacturing site in accordance with the provision of Chapter 4 (limited to Articles 78 and 79 for the manufacturing site that conducts only the labelling, packaging or storing operations) in addition to the provision of Chapter 2. The manufacturer, etc. of medical devices or in-vitro diagnostic reagents shall conduct the manufacturing control and quality control of the products in the manufacturing site specified in Article 96 of Enforcement Regulations in accordance with the provision of Chapter 2 or the provision of Chapter 2 applied mutatis mutandis under Chapter 5, with the proviso that the manufacturing control and quality control in the manufacturing site of the labelling, etc.-category medical device manufacturer, etc. or the labelling, etc.-category in-vitro diagnostic reagents manufacturer, etc. are conducted in accordance with the provision of Chapter 3 or the provision of Chapter 3 applied mutatis mutandis under Chapter 5, instead of the provision of Chapter 2 or the provision of Chapter 2 applied mutatis mutandis under Chapter 5. In addition, the manufacturing control and quality control of the products in the manufacturing site of the biological-origin medical device, etc. manufacturer, etc. shall be conducted in accordance with the provision of Chapter 4 (limited to Articles 78 and 79 for the manufacturing site that conducts only the labelling, packaging or storing operations) in addition to the provision of Chapter 2. The manufacturer of the products concerned with the medical devices or drugs (limited to in-vitro diagnostic reagents, and referred to as such in this Paragraph 4) for the export specified in Paragraph 1 of Article 80 of Law shall conduct the manufacturing control and quality control of the products in the manufacturing site in accordance with the provision of Chapter 2 or the provision of Chapter 2 applied mutatis mutandis under Chapter 5, with the proviso that the labelling, etc.-category medical device manufacturer or the labelling, etc.-category in-vitro diagnostic reagents manufacturer conducts the manufacturing control and quality control of the products concerned with the medical devices or in-vitro diagnostic reagents for the export in accordance with the provision of Chapter 3 or the provision of Chapter 3 applied mutatis mutandis under Chapter 5, instead of the provision of Chapter 2 or the provision of Chapter 2 applied mutatis mutandis under Chapter 5. In addition, the manufacturing control and quality control of the products concerned with the medical devices for the export specified in same Article of Law in the manufacturing site of the biological-origin medical devices, etc. manufacturer shall be conducted in accordance with the provision of Chapter 4 (limited to Articles 78 and 79 for the manufacturing site that conducts only the labelling, packaging or storing operations) in addition to the provision of Chapter 2. | The scope sections of each document set out their objectives. For the ISO/DIS, the objective is to harmonize regulation around the world. As a result FDA made an attempt during the revision of the GMP’s, while the 1996 version of ISO 13485 was being developed, to incorporate the requirements that were included in that version of the Standard. While the agency could not revise the format of the regulation to follow that of the Standard, the requirements were included. The purpose of the JGMP is to align the Quality control and the manufacturing control to the QSR and 13485. This section of the QSR and the JGMP contains some additional regulatory issues that are not appropriate for the Standard. |
| 1.2 Application All requirements of this International Standard are specific to organizations providing medical devices, regardless of the type or size of the organization. If regulatory requirements permit exclusions of design and development controls (see 7.3), this can be used as a justification for their exclusion from the quality management system. These regulations can provide alternative arrangements that are to be addressed in the quality management system. It is the responsibility of the organization to ensure that claims of conformity with this International Standard reflect exclusion of design and development controls [see 4.2.2 a) and 7.3]. If any requirement(s) in Clause 7 of this International Standard is (are) not applicable due to the nature of the medical device(s) for which the quality management system is applied, the organization does not need to include such a requirement(s) in its quality management system [see 4.2.2 a)]. The processes required by this International Standard, which are applicable to the medical device(s), but which are not performed by the organization, are the responsibility of the organization and are accounted for in the organization’s quality management system [see 4.1 a)]. In this International Standard the terms “if appropriate” and “where appropriate” are used several times. When a requirement is qualified by either of these phrases, it is deemed to be “appropriate” unless the organization can document a justification otherwise. A requirement is considered “appropriate” if it is necessary in order for the product to meet specified requirements, and/or the organization to carry out corrective action. | 820.1 Scope (a) Applicability. (1) Current good manufacturing practice (CGMP) requirements are set forth in this quality system regulation. The requirements in this part govern the methods used in, and the facilities and controls used for, the design, manufacture, packaging, labeling, storage, installation, and servicing of all finished devices intended for human use. The requirements in this part are intended to ensure that finished devices will be safe and effective and otherwise in compliance with the Federal Food, Drug, and Cosmetic Act (the act). This part establishes basic requirements applicable to manufacturers of finished medical devices. If a manufacturer engages in only some operations subject to the requirements in this part, and not in others, that manufacturer need only comply with those requirements applicable to the operations in which it is engaged. With respect to class I devices, design controls apply only to those devices listed in Sec. 820.30(a)(2). This regulation does not apply to manufacturers of components or parts of finished devices, but such manufacturers are encouraged to use appropriate provisions of this regulation as guidance. Manufacturers of human blood and blood components are not subject to this part, but are subject to part 606 of this chapter. (2) The provisions of this part shall be applicable to any finished device as defined in this part, intended for human use that is manufactured, imported, or offered for import in any State or Territory of the United States, the District of Columbia, or the Commonwealth of Puerto Rico. (3) In this regulation the term ``where appropriate'' is used several times. When a requirement is qualified by ``where appropriate,'' it is deemed to be ``appropriate'' unless the manufacturer can document justification otherwise. A requirement is ``appropriate'' if nonimplementation could reasonably be expected to result in the product not meeting its specified requirements or the manufacturer not being able to carry out any necessary corrective action. (b) Limitations. The quality system regulation in this part supplements regulations in other parts of this chapter except where explicitly stated otherwise. In the event that it is impossible to comply with all applicable regulations, both in this part and in other parts of this chapter, the regulations specifically applicable to the device in question shall supersede any other generally applicable requirements. | Chapter 2 Manufacturing Control and Quality Control in Manufacturing Sites of Medical Device Manufacturers, etc. Section 1 General Rules Application Article 4 The provisions of Articles 30 to 36 shall not apply to the products concerned 6 with the medical devices other than both the specially designated medical devices specified in Paragraph 1 of Article 77-5 of Law and the medical devices designated by Minister of Health, Labour and Welfare as those of which design and development (hereinafter referred to as “design and development”) are necessary to be controlled to ensure that the manufacturing control and quality control are conducted appropriately. The manufacturer, etc., in case where any provisions of Section 5 of this Chapter are not applicable to their products due to the nature of the medical devices concerned with the products, may not apply such provisions to their quality management system. The manufacturer, etc. shall, in case where either of the provisions specified in preceding two Paragraphs 1 and 2 is applied to, describe the details of such application in the documents which specify the quality management system concerned with the manufacturing site (hereinafter referred to as “quality manual”). | Basically, the applicability guidance for the three documents is the same. In essence, they allow for the exclusion from the QMS requirements associated with activities not performed by the organization. The standard and the JGMP explicitly limits those exclusions to those associated with product realization. Because of this approach, it will be necessary for registrars to explain in detail the scope of any certificates of compliance with ISO 13485:2003. They will have to spell out clearly any exclusion. |
| 2 Normative reference The following referenced documents are indispensable for the application of this document. For dated references, only the edition cited applies. For undated references, the latest edition of the referenced document (including any amendments) applies. ISO 9000:2000, Quality management systems — Fundamentals and vocabulary | N/A | N/A | N/A |
| 3 Terms and definitions For the purposes of this document, the terms and definitions given in ISO 9000 apply, together with the following. The following terms, used in this edition of ISO 13485 to describe the supply chain, have been changed to reflect the vocabulary currently used: Supplier ---> organization ---> customer The term “organization” replaces the term “supplier” used in ISO 13485:1996, and refers to the unit to which this International Standard applies. Also, the term “supplier” now replaces the term “subcontractor”. Throughout the text of this International Standard, wherever the term “product” occurs, it can also mean “service”. Wherever requirements are specified as applying to “medical devices”, the requirements apply equally to related services as supplied by the organization. The following definitions should be regarded as generic, as definitions provided in national regulations can differ slightly and take precedence. | 820.3 Definitions. (a) Act means the Federal Food, Drug, and Cosmetic Act, as amended (secs. 201-903, 52 Stat. 1040 et seq., as amended (21 U.S.C. 321-394)). All definitions in section 201 of the act shall apply to the regulations in this part. | Definitions Article 2 “Process agent” throughout this Ministerial Ordinance means the object that is used for the intermediate products in the manufacturing process (excluding those that constitute parts of the products). “Lot” throughout this Ministerial Ordinance means a grouping of the products, process agents or constituent parts, etc. (hereinafter collectively referred to as “products, etc.”) that are manufactured so as to have a uniform quality in a series of the manufacturing process for a certain manufacturing period. “Batch for testing” throughout this Ministerial Ordinance means a lot or another grouping equivalent thereto of the products with a uniform quality. “Controlled unit” throughout this Ministerial Ordinance means a grouping of the packaging and labelling materials that have been verified to be same. “Input” throughout this Ministerial Ordinance means the information, etc. necessary for the manufacturing control and quality control when conducting a certain process. “Output” throughout this Ministerial Ordinance means the information, etc., that is obtained as a result of a certain process conducted. “Quality management system” throughout this Ministerial Ordinance means the management system that the manufacturer and the foreign manufacturer specified in Paragraph 1 of Article 13-3 of Law (hereinafter simply referred to as “foreign manufacturer”) (hereinafter collectively referred to as “manufacturer, etc.”) direct and control the manufacturing site with regard to quality. “Review” throughout this Ministerial Ordinance means the activity undertaken to determine the suitability, adequacy and effectiveness of the subject matter to achieve the established objectives. “Resource” throughout this Ministerial Ordinance means the personal knowledge and skills and technology, facilities and other resource that is utilised for the operations of the manufacturing site.4 “Infrastructure” throughout this Ministerial Ordinance means the system of the facilities, equipment and services that are necessary for the operations of the manufacturing site. “Traceability” throughout this Ministerial Ordinance means the ability to trace the history, application or location of that which is under consideration. | The Standard spells out clearly the new meanings of the words “supplier” and “organization” The JGMP has more definitions than both of the others, but they relate to pharmaceuticals, and are not part of this comparison. |
| 3.1 Active implantable medical device Active medical device which is intended to be totally or partially introduced, surgically or medically, into the human body or by medical intervention into a natural orifice, and which is intended to remain after the procedure | No definition | No definition | I was not able to find the regulatory definitions of these items, even though I’m sure the definitions in the Standard were taken from a regulatory source. For all intents and purposes I believe that the definitions in 13485 is applicable in all three instances. |
| 3.2 Active medical device Medical device relying for its functioning on a source of electrical energy or any source of power other than that directly generated by the human body or gravity | No definition | No definition |  |
| 3.3 Advisory notice Notice issued by the organization, subsequent to delivery of the medical device, to provide supplementary information and/or to advise what action should be taken in the use of a medical device, the modification of a medical device, the return of the medical device to the organization that supplied it, or the destruction of a medical device NOTE Issue of an advisory notice might be required to comply with national or regional regulations. | No definition | “Advisory notice” throughout this Ministerial Ordinance means the document that is issued subsequent to the delivery of the products to provide supplement any information or to advise what actions should be taken in use, modification, return or destruction of the products. |  |
| 3.4 Customer complaint Written, electronic or oral communication that alleges deficiencies related to the identity, quality, durability, reliability, safety or performance of a medical device that has been placed on the market | (b) Complaint means any written, electronic, or oral communication that alleges deficiencies related to the identity, quality, durability, reliability, safety, effectiveness, or performance of a device after it is released for distribution. | No definition | The regulatory definition is a bit broader than ISO’s, as it covers products that have not, as yet, been placed on the market, but have been released for distribution. This would mean that a lot of product that has been released for distribution could be included in the activities associated with a customer complaint, even though, no part of the lot has not reached the customer. The JGMP does not define complaint, but I presume they accept the QSR definition. |
| 3.5 Implantable medical device Medical device intended to be totally or partially introduced into the human body or a natural orifice, or to replace an epithelial surface or the surface of the eye, by surgical intervention, and which is intended to remain after the procedure for at least 30 days, and which can only be removed by medical or surgical intervention NOTE This definition applies to implantable medical devices other than active implantable medical devices. | [21CFR §812.3(d) Implant means a device that is placed into a surgically or naturally formed cavity of the human body if it is intended to remain there for a period of 30 days or more. FDA man, in order to protect public health, determine that devices placed in subjects for shorter periods are also “implants” for purposes of this part.] | No definition | The definition in the Standard includes explicitly eye implants. I am not sure how FDA treats these products. The JGMP does not specify these, as they are specified in the PAL. |
| No definition | (c) Component means any raw material, substance, piece, part, software, firmware, labeling, or assembly, which is intended to be included as part of the finished, packaged, and labeled device. | “Constituent parts, etc.” throughout this Ministerial Ordinance mean the parts, assemblies (limited to those which are used in the products), raw materials, materials, containers, wrappers, labelling (including the package inserts, and hereinafter referred to as such), etc. used in the manufacturing process that constitute parts of the products, as well as the software of the products. | The requirements are consistent. The definition for ISO is found in ISO 9001. |
| No definition | (d) Control number means any distinctive symbols, such as a distinctive combination of letters or numbers, or both, from which the history of the manufacturing, packaging, labeling, and distribution of a unit, lot, or batch of finished devices can be determined. | No definition |  |
| No definition | (e) Design history file (DHF) means a compilation of records which describes the design history of a finished device. | No definition |  |
| 3.6 Labelling Written, printed or graphic matter affixed to a medical device or any of its containers or wrappers, or accompanying a medical device, related to identification, technical description, and use of the medical device, but excluding shipping documents NOTE Some regional and national regulations refer to “labelling” as “information supplied by the manufacturer.” | [Federal Food, Drug, and Cosmetics Act, Section 201: (m) The term “labeling” means all labels and other written, printed, or graphic matter (1) upon any article or any of its containers or wrappers or (2) accompanying such article. (k) The “label” means a display of written, printed, or graphic matter upon the immediate container of any article; and a requirement made by or under authority of this Act that any word, statement, or other information appear on the label shall not be considered to be complied with unless such word, statement, or other information also appears on the outside container or wrapper, if any there be, of the retail package or such article, or is easily legible through the outside container or wrapper.] | “Packaging and labelling material” throughout this Ministerial Ordinance means the container, wrapper and labelling among the constituent parts, etc. | The definition in the regulation is a bit more detailed, but the definition in the Standard should cover all that is covered in the definitions of “label” and “labelling” in the regulation and the JGMP. |
| [7.3.2 Design and development inputs Inputs relating to product requirements shall be determined and records maintained (see 4.2.4). These inputs shall include a) Functional, performance and safety requirements, according to the intended use, b) Applicable statutory and regulatory requirements, c) Where applicable, information derived from previous similar designs, d) Other requirements essential for design and development, and e) Output(s) of risk management (see 7.1). These inputs shall be reviewed for adequacy and approved. Requirements shall be complete, unambiguous and not in conflict with each other. | (f) Design input means the physical and performance requirements of a device that are used as a basis for device design. | No definition | The definition of “design and development inputs” in the Standard is actually incorporated into the requirements section. It is also a bit more explicit. |
| 7.3.3 Design and development outputs The outputs of design and development shall be provided in a form that enables verification against the design and development input and shall be approved prior to release. Design and development outputs shall a) Meet the input requirements for design and development, b) Provide appropriate information for purchasing, production and for service provision, c) Contain or reference product acceptance criteria, and d) Specify the characteristics of the product that are essential for its safe and proper use. Records of the design and development outputs shall be maintained (see 4.2.4). NOTE Records of design and development outputs can include specifications, manufacturing procedures, engineering drawings, and engineering or research logbooks. | (g) Design output means the results of a design effort at each design phase and at the end of the total design effort. The finished design output is the basis for the device master record. The total finished design output consists of the device, its packaging and labeling, and the device master record. | No definition | The definition of “design and development outputs” in the Standard is actually incorporated into the requirements section. It is also a bit more explicit. While the regulation doesn’t explicitly include examples of design output, it is clear that FDA considers items like the product and component specifications, manufacturing procedures, engineering drawings, and logbooks are part of the design output. |
| 3.7 Medical device Any instrument, apparatus, implement, machine, appliance, implant, in vitro reagent or calibrator, software, material or other similar or related article, intended by the manufacturer to be used, alone or in combination, for human beings for one or more of the specific purpose(s) of diagnosis, prevention, monitoring, treatment or alleviation of disease, diagnosis, monitoring, treatment, alleviation of or compensation for an injury, investigation, replacement, modification, or support of the anatomy or of a physiological process, supporting or sustaining life, control of conception, disinfection of medical devices, providing information for medical purposes by means of in vitro examination of specimens derived from the human body, and which does not achieve its primary intended action in or on the human body by pharmacological, immunological or metabolic means, but which may be assisted in its function by such means. NOTE This definition has been developed by the Global Harmonization Task Force (GHTF). See bibliographic reference [15]. | [Federal Food, Drug, and Cosmetics Act, Section 201: (h) The term “device (…) means an instrument, apparatus, implement, machine, contrivance, implant, in vitro reagent, or other similar or related article, including any component, part, or accessory, which is Recognized in the official National Formulary, or the United States Pharmacopoeia, or any supplement to them, Intended for use in the diagnosis of disease or other conditions, or in the cure, mitigation, treatment, or prevention of disease, in man or other animals, or Intended to affect the structure or any function of the body of man or other animals, and which does not achieve its primary intended purposes through chemical action within or on the body of man or other animals and which is not dependent upon being metabolized for the achievement of its primary intended purposes.] | No definition | No significant difference. |
| 7.3.4 Design and development review At suitable stages, systematic reviews of design and development shall be performed in accordance with planned arrangements (see 7.3.1) a) To evaluate the ability of the results of design and development to meet requirements, and b) To identify any problems and propose necessary actions. Participants in such reviews shall include representatives of functions concerned with the design and development stage(s) being reviewed, as well as other specialist personnel (see 5.5.1 and 6.2.1). Records of the results of the reviews and any necessary actions shall be maintained (see 4.2.4). | (h) Design review means a documented, comprehensive, systematic examination of a design to evaluate the adequacy of the design requirements, to evaluate the capability of the design to meet these requirements, and to identify problems. | No definition | The definition of “design and development review” in the Standard is actually incorporated into the requirements section. Otherwise there is no significant difference. |
| 3.8 Sterile medical device Category of medical device intended to meet the requirements for sterility NOTE The requirements for sterility of a medical device might be subject to national or regional regulations or standards |  | “Sterile medical device” throughout this Ministerial Ordinance means a medical device that is sterilised in the manufacturing process. |  |
| No definition | (i) Device history record (DHR) means a compilation of records containing the production history of a finished device. (j) Device master record (DMR) means a compilation of records containing the procedures and specifications for a finished device. (k) Establish means define, document (in writing or electronically), and implement. (l) Finished device means any device or accessory to any device that is suitable for use or capable of functioning, whether or not it is packaged, labeled, or sterilized. (m) Lot or batch means one or more components or finished devices that consist of a single type, model, class, size, composition, or software version that are manufactured under essentially the same conditions and that are intended to have uniform characteristics and quality within specified limits. | No definition |  |
| [ISO 9000:2000, 3.2.7 top management – person or group of people who directs and controls an organization (3.3.1) at the highest level.] | (n) Management with executive responsibility means those senior employees of a manufacturer who have the authority to establish or make changes to the manufacturer's quality policy and quality system. | “Top management” throughout this Ministerial Ordinance means the person or the group of people, including the executive officer who conducts the duties concerned with the manufacturing site, who directs and controls the manufacturing site. | The intent of the definitions is the same. |
| No definition | (o) Manufacturer means any person who designs, manufactures, fabricates, assembles, or processes a finished device. Manufacturer includes but is not limited to those who perform the functions of contract sterilization, installation, relabelling, remanufacturing, repacking, or specification development, and initial distributors of foreign entities performing these functions. (p) Manufacturing material means any material or substance used in or used to facilitate the manufacturing process, a concomitant constituent, or a byproduct constituent produced during the manufacturing process, which is present in or on the finished device as a residue or impurity not by design or intent of the manufacturer. | The definition is included in the scope. | The requirements are consistent |
| [ISO 9000:2000, 3.6.2 nonconformity – non-fulfilment of a requirement (3.1.2)] | (q) Nonconformity means the nonfulfillment of a specified requirement. | No definition | Identical |
| [ISO 9000:2000, 3.4.2 product – the result of a process (3.4.1) NOTE 1 There are four generic product categories, as follows: Services (e.g., transport); Software (e.g., computer program, dictionary); Hardware (e.g., engine mechanical part); Processed materials (e.g., lubricant). Many products comprise elements belonging to different generic product categories. Whether the product is then called service, software, hardware or processed material depends on the dominant element. For example, the offered product “automobile” consists of hardware (e.g., tires), processed material (e.g., fuel, cooling system liquid), software (e.g., engine control software, driver’s manual), and service (e.g., operating explanations given by the salesman). …] | (r) Product means components, manufacturing materials, in- process devices, finished devices, and returned devices. | “Product” throughout this Ministerial Ordinance means the object (including those which have undergone the intermediate process and need to undergo subsequent process to be the (final) products (hereinafter referred to as “intermediate products”)) that has undergone the manufacturing process in the manufacturing site. | The definition in the Standard more clearly reflects the process approach of this document. The definition in the Standard is a bit more detailed. The QSR and the JGMP is consistent. |
| [ISO 9000:2000, 3.1.1 quality – degree to which a set of inherent characteristics (3.5.1) fulfils requirements (3.1.2)] | (s) Quality means the totality of features and characteristics that bear on the ability of a device to satisfy fitness-for-use, including safety and performance. | No definition | No significant difference. The definition in the regulation reflects the objective of the regulation, that is the assurance of product safety and effectiveness. |
| [ISO 90002000, 3.9.1 audit – systematic, independent and documented process (3.4.1) for obtaining audit evidence (3.9.4) and evaluating it objectively to determine the extent to which audit criteria (3.9.3) are fulfilled …] | (t) Quality audit means a systematic, independent examination of a manufacturer's quality system that is performed at defined intervals and at sufficient frequency to determine whether both quality system activities and the results of such activities comply with quality system procedures, that these procedures are implemented effectively, and that these procedures are suitable to achieve quality system objectives. | No definition | The definition in the regulation imposes the requirement related to defined intervals, and includes activities that occur after the audit is actually performed. |
| [ISO 9000:2000, 3.2.4 quality policy – overall intentions and direction of an organization (3.3.1) related to quality (3.1.1) as formally expressed by top management ((3.2.7)] | (u) Quality policy means the overall intentions and direction of an organization with respect to quality, as established by management with executive responsibility. | “Quality policy” throughout this Ministerial Ordinance means the overall intentions and direction of the manufacturing site related to the quality as formally provided and expressed by the top management in order to ensure the quality of the products. | No significant difference. |
| [ISO 9000:2000, 3.2.2 system – set of interrelated or interacting elements] | (v) Quality system means the organizational structure, responsibilities, procedures, processes, and resources for implementing quality management. | No definition | The definition in the regulation is more instructive, but is consistent with the intent of the Standard. |
|  | (w) Remanufacturer means any person who processes, conditions, renovates, repackages, restores, or does any other act to a finished device that significantly changes the finished device's performance or safety specifications, or intended use. | No definition |  |
| [ISO 9000:2000, 3.6.7 rework – action on a nonconforming product (3.4.2) to make it conform to the requirements (3.1.2) NOTE Unlike rework, repair (3.6.9) can affect or change parts of the nonconforming product.] | (x) Rework means action taken on a nonconforming product so that it will fulfill the specified DMR requirements before it is released for distribution. | No definition | No significant difference. |
| [ISO 9000:2000, 3.7.3 specification – document (3.7.2) stating requirements (3.1.2) NOTE A specification can be related to activities (e.g., procedure document, process specification and test specification), or products (3.4.2) (e.g., product specification, performance specification and drawing)] | (y) Specification means any requirement with which a product, process, service, or other activity must conform. | No definition | The regulation requires documentation of specifications. |
| [ISO 9000:2000, 3.8.5 validation – confirmation, through the provision of objective evidence (3.8.1), that the requirements (3.1.2) for a specific intended use or application have been fulfilled. NOTE 1 The term “validated” is used to designate the corresponding status NOTE 2 The use conditions for validation can be real or simulated] | (z) Validation means confirmation by examination and provision of objective evidence that the particular requirements for a specific intended use can be consistently fulfilled. (1) Process validation means establishing by objective evidence that a process consistently produces a result or product meeting its predetermined specifications. (2) Design validation means establishing by objective evidence that device specifications conform with user needs and intended use(s). | “Validation” throughout this Ministerial Ordinance means to verify and document that the buildings and facilities of the manufacturing site, procedures, processes and other procedures of the manufacturing control and quality control (hereinafter referred to as “manufacturing procedure, etc.”) provide the anticipated results.3 | No significant difference, but neither definition is very informative. However, JGMP specifically mentions buildings and facilities. The regulatory definition is more detailed in that it breaks out and defines “process validation” and “design validation”. Normally, validation is performed on the final product or process for manufacturing, monitoring, testing, and supporting the product. |
| [ISO 9000:2000, 3.8.4 verification – confirmation, through the provision of objective evidence (3.8.1) that specified requirements (3.1.2) have been fulfilled NOTE 1 The term “verified” is used to designate the corresponding status. NOTE 2 Confirmation can comprise activities such as Performing alternative calculations Comparing a new design specification (3.7.3) with a similar proven design specification, Undertaking tests (3.8.3) and demonstrations, and Reviewing documents prior to issue.] | (aa) Verification means confirmation by examination and provision of objective evidence that specified requirements has been fulfilled. | No definition | No significant difference. The Standard is more detailed. |
| 4 Quality management system 4.1 General requirements The organization shall establish, document, implement and maintain a quality management system and maintain its effectiveness in accordance with the requirements of this International Standard. The organization shall a) Identify the processes needed for the quality management system and their application throughout the organization (see 1.2), b) Determine the sequence and interaction of these processes, c) Determine criteria and methods needed to ensure that both the operation and control of these processes are effective, d) Ensure the availability of resources and information necessary to support the operation and monitoring of these processes, e) Monitor, measure and analyse these processes, and f) Implement actions necessary to achieve planned results and maintain the effectiveness of these processes. These processes shall be managed by the organization in accordance with the requirements of this International Standard. Where an organization chooses to outsource any process that affects product conformity with requirements, the organization shall ensure control over such processes. Control of such outsourced processes shall be identified within the quality management system (see 8.5.1). NOTE Processes needed for the quality management system referred to above should include processes for management activities, provision of resources, product realization and measurement. | 820.5 Quality system. Each manufacturer shall establish and maintain a quality system that is appropriate for the specific medical device(s) designed or manufactured, and that meets the requirements of this part. | Section 2 Quality Management System General Requirements for Quality Management System Article 5 The manufacturer, etc. shall establish, document, implement and maintain a quality management system and maintain its effectiveness in accordance with the provision of this Chapter. The manufacturer, etc. shall conduct the following duties. To identify the processes needed for the quality management system (including the results to be achieved by the processes) (hereinafter simply referred to as “processes” in this Chapter) and their application throughout the manufacturing site, To determine the sequence and interaction of the processes, To determine the criteria and methods needed to ensure that both the operation and control of the processes are effective, To ensure the availability of the resources and information necessary to support the operation and monitoring of the processes, To monitor, measure and analyse these processes, and To implement the actions necessary to achieve the planned results specified in preceding Item (1) and maintain the effectiveness of these processes. The manufacturer, etc. shall manage the processes in accordance with the provision of this Chapter. 4. Where the manufacturer, etc. choose to outsource any process (excluding those which are subject to the authorisation specified in Paragraph 1 of Article 13 of Law or the recognition specified in Paragraph 1 of Article 13-3 of Law) that affects the conformity7 with the product requirements (including the laws, orders and ordinances which are relevant to the pharmaceutical affairs or the orders or official actions based on the laws, orders and ordinances (hereinafter simply referred to as “regulatory requirements” in this Chapter)) (hereinafter referred to as “product requirements” in this Chapter), the manufacturer, etc. shall ensure the control over such processes. The manufacturer, etc. shall ensure that the control of such outsourced processes is identified within the quality management system. | No significant differences |
| 4.2 Documentation requirements 4.2.1 General The quality management system documentation shall include a) Documented statements of a quality policy and quality objectives, b) A quality manual, c) Documented procedures required by this International Standard, d) Documents needed by the organization to ensure the effective planning, operation and control of its processes, e) Records required by this International Standard (see 4.2.4), and f) Any other documentation specified by national or regional regulations. Where this International Standard specifies that a requirement, procedure, activity or special arrangement be “documented”, it shall, in addition, be implemented and maintained. For each type or model of medical device, the organization shall establish and maintain a file either containing or identifying documents defining product specifications and quality management system requirements (see 4.2.3). These documents shall define the complete manufacturing process and, if applicable, installation and servicing. NOTE 1 The extent of the quality management system documentation can differ from one organization to another due to a) The size of the organization and type of activities, b) The complexity of processes and their interactions, and c) The competence of personnel. NOTE 2 The documentation can be in any form or type of medium. |  | Documentation of Quality Management System Article 6 The manufacturer, etc. shall ensure that the quality management system documentation includes the following documents and that the documented requirement, procedure, activity or special arrangement is implemented and maintained. The documented statements of the quality policy and quality objectives, The quality manual, The documents needed by the manufacturer, etc. to ensure the effective planning, operation and control of their processes in the manufacturing site, The documented procedure and records specified in this Chapter, and Any other documentation specified by the laws, orders and ordinances related to the pharmaceutical affairs. 2. For each type or model of the medical devices, the manufacturer, etc. shall establish and maintain the document defining the product specifications and quality management system requirements (hereinafter referred to as “Seihin Hyojun Sho” in this Chapter), or a file identifying Seihin Hyojun Sho. The manufacturer, etc. shall ensure that Seihin Hyojun Sho defines the complete manufacturing process concerned with the products of the manufacturing site and, if applicable, the installation specified in Paragraph 1 of Article 42 and the servicing specified in Paragraph 1 of Article 43. | The ISO13485 standard and the JGMP lists the required quality management system documentation in this clause, while the QSR indicates the documentation requirements in the various sections throughout the regulation. There is no significant difference in the documentation requirements for the three documents, except as indicated below. The QSR implicitly recognizes that the extent of quality management system documentation will reflect the size and complexity of the organization. It also implicitly recognizes that some documentation may not be needed due to the expertise (either through training, education, or experience) of the personnel. The JGMP has a unique requirement to define the processes. This is consistent with the intent of 13485. This text is included in ISO 13485:2003 in order to accommodate the definition of “establish” in the QS Reg. |
| 4.2.2 Quality manual The organization shall establish and maintain a quality manual that includes a) The scope of the quality management system, including details of and justification for any exclusion and/or non-application (see 1.2), b) The documented procedures established for the quality management system, or reference to them, and c) A description of the interaction between the processes of the quality management system. The quality manual shall outline the structure of the documentation used in the quality management system. | No definition | Quality Manual Article 7 1. The manufacturer, etc. shall establish and maintain the quality manual that includes the following items. The scope of the quality management system, including the details of and justification for any exclusion and/or non-application, The documented procedure established for the quality management system, or reference to them, and The description of the interaction between the processes of the quality management system.8 2. The manufacturer, etc. shall ensure that the quality manual outlines the structure of the documentation used in the quality management system. | The QSR has no requirement for a Quality Manual. Such a manual would still be helpful in explaining the nature and extent of the quality management system to an FDA investigator during an inspection. It would also be useful in the training of personnel with regard to the quality management system of the organization and their place within that system. The Quality Manual could be used as the repository of some of the individual quality management system documentation required by the QSR (e.g., the organizational structure and interrelationships, the highest level procedures in a small organization dealing with items like document control, records keeping, training) |
| 4.2.3 Control of documents Documents required by the quality management system shall be controlled. Records are a special type of document and shall be controlled according to the requirements given in 4.2.4. A documented procedure shall be established to define the controls needed a) To review and approve documents for adequacy prior to issue, b) To review and update as necessary and re-approve documents, c) To ensure that changes and the current revision status of documents are identified, d) To ensure that relevant versions of applicable documents are available at points of use, e) To ensure that documents remain legible and readily identifiable, f) To ensure that documents of external origin are identified and their distribution controlled, and g) To prevent the unintended use of obsolete documents, and to apply suitable identification to them if they are retained for any purpose. The organization shall ensure that changes to documents are reviewed and approved either by the original approving function or another designated function, which has access to pertinent background information upon which to base its decisions. The organization shall define the period for which at least one copy of obsolete controlled documents shall be retained. This period shall ensure that documents to which medical devices have been manufactured and tested are available for at least the lifetime of the medical device as defined by the organization, but not less than the retention period of any resulting record (see 4.2.4), or as specified by relevant regulatory requirements. | 820.40 Document controls. Each manufacturer shall establish and maintain procedures to control all documents that are required by this part. The procedures shall provide for the following: (a) Document approval and distribution. Each manufacturer shall designate an individual(s) to review for adequacy and approve prior to issuance all documents established to meet the requirements of this part. The approval, including the date and signature of the individual(s) approving the document, shall be documented. Documents established to meet the requirements of this part shall be available at all locations for which they are designated, used, or otherwise necessary, and all obsolete documents shall be promptly removed from all points of use or otherwise prevented from unintended use. (b) Document changes. Changes to documents shall be reviewed and approved by an individual(s) in the same function or organization that performed the original review and approval, unless specifically designated otherwise. Approved changes shall be communicated to the appropriate personnel in a timely manner. Each manufacturer shall maintain records of changes to documents. Change records shall include a description of the change, identification of the affected documents, the signature of the approving individual(s), the approval date, and when the change becomes effective. | Control of Documents Article 8 The manufacturer, etc. shall ensure that the documents (excluding the records), including those specified in preceding two Articles and other Articles in this Chapter, required by the quality management system (hereinafter referred to as “QMS documents”) are controlled. The manufacturer, etc. shall ensure that the documented procedure is established to define the controls needed for the following duties. To review and approve the QMS documents for adequacy prior to the issue, To review and update as necessary and re-approve the QMS documents, To ensure that the changes and current revision status of the QMS documents are identified, To ensure that the relevant versions of the applicable QMS documents are available at the points of the use, To ensure that the QMS documents remain legible and readily identifiable, To ensure that the QMS documents of external origin are identified and their distribution is controlled, and To prevent the unintended use of the obsolete QMS documents, and to apply suitable identification to them if they are retained for any purpose. The manufacturer, etc. shall ensure that the changes to the QMS documents are reviewed and approved either by the original approving department or another designated department which has access to pertinent background information upon which to base its decisions. The manufacturer, etc. shall retain at least one copy of the obsolete QMS documents for the following period (5 years for the QMS documents concerned with the training) from the date of obsolescence, with the proviso that this provision shall not apply to the documents used for the manufacturing or testing of the products in case where they are maintained to be available for the retention period of the records of the products specified in next Article. 15 years for the products concerned with the specially designated maintenance-control-required medical devices specified in Paragraph 8 of Article 2 of Law (1 year plus the shelf life or the period until the expiry date (hereinafter simply referred to as “shelf life”), for the products concerned with the medical devices of which shelf life shall be displayed and of which shelf life plus 1 year exceeds 15 years), or9 5 years for the products concerned with the medical devices other than the specially designated maintenance-control-required medical devices (1 year plus the shelf life, for the products of which shelf life shall be displayed and of which shelf life plus 1 year exceeds 5 years). | The requirements are essentially the same, except that the QSR has the specific requirement to communicate changes to documents to the affected personnel. ISO/DIS 13485:2003 requires the organization to define a retention period for obsoleted documents. The JGMP specifies the retention period depending on the type of medical device. |
| 4.2.4 Control of records Records shall be established and maintained to provide evidence of conformity to requirements and of the effective operation of the quality management system. Records shall remain legible, readily identifiable and retrievable. A documented procedure shall be established to define the controls needed for the identification, storage, protection, retrieval, retention time and disposition of records. The organization shall retain the records for a period of time at least equivalent to the lifetime of the medical device as defined by the organization, but not less than two years from the date of product release by the organization or as specified by relevant regulatory requirements. | 820.180 General requirements. All records required by this part shall be maintained at the manufacturing establishment or other location that is reasonably accessible to responsible officials of the manufacturer and to employees of FDA designated to perform inspections. Such records, including those not stored at the inspected establishment, shall be made readily available for review and copying by FDA employee(s). Such records shall be legible and shall be stored to minimize deterioration and to prevent loss. Those records stored in automated data processing systems shall be backed up. (a) Confidentiality. Records deemed confidential by the manufacturer may be marked to aid FDA in determining whether information may be disclosed under the public information regulation in part 20 of this chapter. (b) Record retention period. All records required by this part shall be retained for a period of time equivalent to the design and expected life of the device, but in no case less than 2 years from the date of release for commercial distribution by the manufacturer. (c) Exceptions. This section does not apply to the reports required by Sec. 820.20(c) Management review, Sec. 820.22 Quality audits, and supplier audit reports used to meet the requirements of Sec. 820.50(a) Evaluation of suppliers, contractors, and consultants, but does apply to procedures established under these provisions. Upon request of a designated employee of FDA, an employee in management with executive responsibility shall certify in writing that the management reviews and quality audits required under this part, and supplier audits where applicable, have been performed and documented, the dates on which they were performed, and that any required corrective action has been undertaken. 820.181 Device master record. Each manufacturer shall maintain device master records (DMR's). Each manufacturer shall ensure that each DMR is prepared and approved in accordance with Sec. 820.40. The DMR for each type of device shall include, or refer to the location of, the following information: (a) Device specifications including appropriate drawings, composition, formulation, component specifications, and software specifications; (b) Production process specifications including the appropriate equipment specifications, production methods, production procedures, and production environment specifications; (c) Quality assurance procedures and specifications including acceptance criteria and the quality assurance equipment to be used; (d) Packaging and labeling specifications, including methods and processes used; and (e) Installation, maintenance, and servicing procedures and methods. 820.184 Device history record. Each manufacturer shall maintain device history records (DHR's). Each manufacturer shall establish and maintain procedures to ensure that DHR's for each batch, lot, or unit are maintained to demonstrate that the device is manufactured in accordance with the DMR and the requirements of this part. The DHR shall include, or refer to the location of, the following information: (a) The dates of manufacture; (b) The quantity manufactured; (c) The quantity released for distribution; (d) The acceptance records which demonstrate the device is manufactured in accordance with the DMR; (e) The primary identification label and labeling used for each production unit; and (f) Any device identification(s) and control number(s) used. 820.186 Quality system record. Each manufacturer shall maintain a quality system record (QSR). The QSR shall include, or refer to the location of, procedures and the documentation of activities required by this part that are not specific to a particular type of device(s), including, but not limited to, the records required by Sec. 820.20. Each manufacturer shall ensure that the QSR is prepared and approved in accordance with Sec. 820.40. | Control of Records Article 9 The manufacturer, etc. shall ensure that the records are established and maintained to provide the evidence of the conformity to the requirements, including those for the records, specified in this Chapter and of the effective operation of the quality management system and that the records remain legible, readily identifiable and retrievable. 2. The manufacturer, etc. shall ensure that the documented procedure is established to define the controls needed for the identification, storage, protection, retrieval, retention time and disposition of the records specified in preceding Paragraph 1. 3. The manufacturer, etc. shall retain the records specified in preceding Paragraph 1 for the following period (5 years for the records concerned with the training) from the date of the establishment. (1) 15 years for the products concerned with the specially designated maintenance-control- required medical devices (1 year plus the shelf life, for the products concerned with the medical devices of which shelf life shall be displayed and of which shelf life plus 1 year exceeds 15 years), or (2) 5 years for the products concerned with the medical devices other than the specially designated maintenance-control-required medical devices (1 year plus the shelf life, for the products concerned with the medical devices of which shelf life shall be displayed and of which shelf life plus 1 year exceeds 5 years). | No significant differences in the general requirements associated with control of records, except that the QSR contains requirements for communications with FDA. The JGMP specifies the retention period depending on the type of medical device. The QSR requires the establishment of a Device Master Record (DMR). The DMR may be a separate file of documents and records, or it may be documents containing references to the various elements of the DMR. ISO/DIS 13485:2003 and the JGMP has no requirement for such a file even though it requires the individual documents and records that would be contained within that file. The QSR requires the establishment of a Device History Record (DHR) for each lot of devices or unit manufactured. The DHR may be a separate file containing the records, or may be document that references the location of these records. ISO/DIS 13485:2003 and JGMP does not require the establishment of such a file, even though it requires the individual records that would be contained with that file. The QSR requires the establishment of a Quality System Record, which may be a separate file containing the required documents or a document referencing the required contents. ISO/DIS 13485:2003 and JGMP does not require the establishment of such a file, even though it does requirement the establishment of the various documents that would be included in that file. |
| 5 Management responsibility 5.1 Management commitment Top management shall provide evidence of its commitment to the development and implementation of the quality management system and maintaining its effectiveness by a) Communicating to the organization the importance of meeting customer as well as statutory and regulatory requirements, b) Establishing the quality policy, c) Ensuring that quality objectives are established, d) Conducting management reviews, and e) Ensuring the availability of resources. NOTE For the purposes of this International Standard, statutory requirements are limited to the safety and performance of the medical device only. | 820.20 Management responsibility. (a) Quality policy. Management with executive responsibility shall establish its policy and objectives for, and commitment to, quality. Management with executive responsibility shall ensure that the quality policy is understood, implemented, and maintained at all levels of the organization. | Section 3 Management Responsibility Management Commitment Article 10 The top management shall provide evidence of its commitment to the development and implementation of the quality management system and maintaining its effectiveness by conducting the following duties. To establish the quality policy, To ensure that the quality objectives are established, To conduct the management reviews specified in Paragraph 1 of Article 18, To ensure the availability of the resources, and To communicate to the manufacturing site the importance of meeting the10 requirements of the marketing authorisation holder and other bodies receiving the products (hereinafter referred to as “customers”) (hereinafter referred to as “customer requirements”) as well as the regulatory requirements. | No significant differences in management responsibilities; the QSR spells some of them out in subsequent sections of the regulation. The JGMP specifically mentions the 10 requirements for the marketing authorization holder. |
| 5.2 Customer focus Top management shall ensure that customer requirements are determined and are met (see 7.2.1 and 8.2.1). |  | Customer Focus Article 11 The top management shall ensure that the customer requirements are determined and are met. | ISO/DIS 13485:2003 and JGMP have a distinct focus on meeting customer requirements in addition to meeting regulatory requirements. The QSR is entirely focused on meeting those requirements that have as their objective the design, manufacture, distribution, and support of safe and effective medical devices. ISO/DIS 13485:2003 will include requirements for determining customer requirements during the entire product realization process, while the QSR will include requirements that identify product and process requirements that are focused on ensuring safe and effective medical devices. |
| 5.3 Quality policy Top management shall ensure that the quality policy a) Is appropriate to the purpose of the organization, b) Includes a commitment to comply with requirements and to maintain the effectiveness of the quality management system, c) Provides a framework for establishing and reviewing quality objectives, d) Is communicated and understood within the organization, and e) Is reviewed for continuing suitability. | 820.20 Management responsibility. (a) Quality policy. Management with executive responsibility shall establish its policy and objectives for, and commitment to, quality. Management with executive responsibility shall ensure that the quality policy is understood, implemented, and maintained at all levels of the organization. | Quality Policy Article 12 The top management shall ensure that the quality policy meets the following requirements. To be appropriate to the purpose of the manufacturer, etc., To include the commitment to comply with the requirements and to maintain the effectiveness of the quality management system, To provide the framework for establishing and reviewing the quality objectives, To be communicated and understood within the manufacturing site, and To be reviewed for continuing suitability. | No significant differences |
| 5.4 Planning 5.4.1 Quality objectives Top management shall ensure that quality objectives, including those needed to meet requirements for product [see 7.1 a)], are established at relevant functions and levels within the organization. The quality objectives shall be measurable and consistent with the quality policy. | 820.20 Management responsibility. (a) Quality policy. Management with executive responsibility shall establish its policy and objectives for, and commitment to, quality. Management with executive responsibility shall ensure that the quality policy is understood, implemented, and maintained at all levels of the organization. | Quality Objectives Article 13 1. The top management shall ensure that the quality objectives, including those needed to meet the requirements for the products, are established at relevant departments within the manufacturing site. 2. The top management shall ensure that the quality objectives are measurable and consistent with the quality policy. | No significant differences |
| 5.4.2 Quality management system planning Top management shall ensure that a) The planning of the quality management system is carried out in order to meet the requirements given in 4.1, as well as the quality objectives, and b) The integrity of the quality management system is maintained when changes to the quality management system are planned and implemented. | 820.5 Quality system. Each manufacturer shall establish and maintain a quality system that is appropriate for the specific medical device(s) designed or manufactured, and that meets the requirements of this part. (d) Quality planning. Each manufacturer shall establish a quality plan, which defines the quality practices, resources, and activities relevant to devices that are designed and manufactured. The manufacturer shall establish how the requirements for quality will be met. (e) Quality system procedures. Each manufacturer shall establish quality system procedures and instructions. An outline of the structure of the documentation used in the quality system shall be established where appropriate. | Quality Management System Planning Article 14 1. The top management shall ensure that the planning of the quality management system is carried out in order to meet the requirements specified in Article 5, as well as the quality objectives. 2. The top management shall ensure that the integrity of the quality management system is maintained when changes to the quality management system are planned and implemented. | The QSR contains the prescriptive requirements for a quality plan and quality system procedures. It is not clear what the FDA is looking for when they ask for a quality plan. It seems to be combination of a high level quality planning document, containing policy and key objectives, with a mandate to drive those objectives down into the organization, and a set of high level procedures that illustrate how that plan will be met. The QSR, JGMP, and ISO/DIS 13485:2003 require the establishment of these kinds of procedures; only the QSR gives them special standing as quality system procedures. |
| 5.5 Responsibility, authority and communication 5.5.1 Responsibility and authority Top management shall ensure that responsibilities and authorities are defined, documented and communicated within the organization. Top management shall establish the interrelation of all personnel who manage, perform and verify work, affecting quality, and shall ensure the independence and authority necessary to perform these tasks. NOTE National or regional regulations might require the nomination of specific persons as responsible for activities related to monitoring experience from the post-production stage and reporting adverse events (see 8.2.1 and 8.5.1). | Management responsibility (b) Organization. Each manufacturer shall establish and maintain an adequate organizational structure to ensure that devices are designed and produced in accordance with the requirements of this part. (1) Responsibility and authority. Each manufacturer shall establish the appropriate responsibility, authority, and interrelation of all personnel who manage, perform, and assess work, affecting quality, and provide the independence and authority necessary to perform these tasks. | Responsibility and Authority Article 15 1. The top management shall ensure that the responsibilities and authorities of the departments and personnel engaged in the duties are defined, documented and communicated within the manufacturing site.11 2. The top management shall establish the interrelation of all personnel who manage, perform and verify the work affecting the quality, and shall ensure the independence and authority necessary to perform these tasks. | No significant differences |
| 5.5.2 Management representative Top management shall appoint a member of management who, irrespective of other responsibilities, shall have responsibility and authority that includes a) Ensuring that processes needed for the quality management system are established, implemented and maintained, b) Reporting to top management on the performance of the quality management system and any need for improvement (see 8.5), and c) Ensuring the promotion of awareness of regulatory and customer requirements throughout the organization. NOTE The responsibility of a management representative can include liaison with external parties on matters relating to the quality management system. | Management responsibility Organization (3) Management representative. Management with executive responsibility shall appoint, and document such appointment of, a member of management who, irrespective of other responsibilities, shall have established authority over and responsibility for: (i) Ensuring that quality system requirements are effectively established and effectively maintained in accordance with this part; and (ii) Reporting on the performance of the quality system to management with executive responsibility for review. | Responsible Engineering Manager Article 16 1. The top management shall ensure that the responsible engineering manager specified in Paragraph 5 of Article 17 of Law and the manager controlling the manufacturing of the biological-origin products (the biological-origin products specified in Paragraph 9 of Article 2 of Law, and hereinafter referred to as such) specified in Paragraph 1 of Article 68-2 of Law (in case of a foreign manufacturer, the person responsible for the manufacturing site which has been recognized in accordance with the provision of Paragraph 1 of Article 13-3 of Law or the person designated beforehand by such a foreign manufacturer) (hereinafter collectively and simply referred to as “responsible engineering manager”) have the responsibility and authority for the following duties. To ensure that the processes needed for the quality management system are established, implemented and maintained, To report to the top management on the performance of the quality management system and any need for improvement, and To ensure the promotion of the awareness of the regulatory and customer requirements throughout the manufacturing site. | No significant differences, except the requirements in ISO/DIS 13485:2003 and JGMP reflect the focus on meeting customer requirements. |
| 5.5.3 Internal communication Top management shall ensure that appropriate communication processes are established within the organization and that communication takes place regarding the effectiveness of the quality management system. | Management responsibility (b) Organization. Each manufacturer shall establish and maintain an adequate organizational structure to ensure that devices are designed and produced in accordance with the requirements of this part. (1) Responsibility and authority. Each manufacturer shall establish the appropriate responsibility, authority, and interrelation of all personnel who manage, perform, and assess work, affecting quality, and provide the independence and authority necessary to perform these tasks. | Internal Communication Article 17 The top management shall ensure that appropriate communication processes are established within the manufacturing site and that communication takes place regarding the effectiveness of the quality management system. | No significant differences; the QSR implicitly requires the necessary communication processes that make for successful interrelationships |
| 5.6 Management review 5.6.1 General Top management shall review the organization’s quality management system, at planned intervals, to ensure its continuing suitability, adequacy and effectiveness. This review shall include assessing opportunities for improvement and the need for changes to the quality management system, including the quality policy and quality objectives. Records from management reviews shall be maintained (see 4.2.4). | Management responsibility (c) Management review. Management with executive responsibility shall review the suitability and effectiveness of the quality system at defined intervals and with sufficient frequency according to established procedures to ensure that the quality system satisfies the requirements of this part and the manufacturer's established quality policy and objectives. The dates and results of quality system reviews shall be documented. | Management Review Article 18 1. The top management shall review the manufacturing site's quality management system, at the interval defined in the plan specified in Paragraph 1 of Article 14, to ensure its continuing suitability, adequacy and effectiveness. This review (hereinafter referred to as “management review”) shall include assessing the opportunities for improvement and the need for changes to the quality management system, including the quality policy and quality objectives. 2. The manufacturer, etc. shall ensure that the records from the management reviews are maintained. | No significant differences. |
| 5.6.2 Review input The input to management review shall include information on a) Results of audits, b) Customer feedback, c) Process performance and product conformity, d) Status of preventive and corrective actions, e) Follow-up actions from previous management reviews, f) Changes that could affect the quality management system, g) Recommendations for improvement, and h) New or revised regulatory requirements |  | Review Input Article 19 The manufacturer, etc. shall ensure that the input to the management review includes the information on the following items.12 (1) The results of the audits, (2) The customer feedback, (3) The process performance and product conformity to the product requirements, (4) The status of the corrective actions (the actions taken to eliminate the cause of the detected non-conformity or other undesirable situation, and hereinafter referred to as such) and the preventive actions (the actions to eliminate the cause of the potential non-conformity or other undesirable potential situation, and hereinafter referred to as such), (5) The follow-up actions from the previous management reviews, (6) The changes that could affect the quality management system, (7) The recommendations for improvement, and (8) New or revised regulatory requirements. | The requirements for review input that are spelled out in ISO/DIS 13485:2003 and JGMP are logical and would be expected by an FDA investigator during an inspection that focused on management responsibilities. |
| 5.6.3 Review output The output from the management review shall include any decisions and actions related to a) Improvements needed to maintain the effectiveness of the quality management system and its processes, b) Improvement of product related to customer requirements, and c) Resource needs. |  | Review Output Article 20 The manufacturer, etc. shall ensure that the output from the management review includes any decisions and actions related to the following items. (1) The improvements needed to maintain the effectiveness of the quality management system and its processes, (2) The improvement of the products related to the customer requirements, and (3) The resource needs to ensure continuing suitability, adequacy and effectiveness of the quality management system. | The requirements for review output that are spelled out in ISO/DIS 13485:2003 are logical and would be expected by an FDA investigator during an inspection that focused on management responsibilities. |
| 6 Resource management 6.1 Provision of resources The organization shall determine and provide the resources needed a) To implement the quality management system and to maintain its effectiveness, and b) To meet regulatory and customer requirements | Management responsibility (b) Organization (2) Resources. Each manufacturer shall provide adequate resources, including the assignment of trained personnel, for management, performance of work, and assessment activities, including internal quality audits, to meet the requirements of this part. | Section 4 Resource Management Provision of Resources Article 21 The manufacturer, etc. shall determine and provide the resources needed for the following duties. (1) To implement the quality management system and to maintain its effectiveness, and (2) To meet the regulatory and customer requirements. | No significant differences, especially since ISO/DIS 13485:2003 and the JGMP include the requirement to meeting regulatory requirements. |
| 6.2 Human resources 6.2.1 General Personnel performing work affecting product quality shall be competent on the basis of appropriate education, training, skills and experience. | 820.20 Management responsibility (b) Organization (2) Resources. Each manufacturer shall provide adequate resources, including the assignment of trained personnel, for management, performance of work, and assessment activities, including internal quality audits, to meet the requirements of this part. 820.25 Personnel. (a) General. Each manufacturer shall have sufficient personnel with the necessary education, background, training, and experience to assure that all activities required by this part are correctly performed. (b) Training. Each manufacturer shall establish procedures for identifying training needs and ensure that all personnel are trained to adequately perform their assigned responsibilities. Training shall be documented. (1) As part of their training, personnel shall be made aware of device defects, which may occur from the improper performance of their specific jobs. (2) Personnel who perform verification and validation activities shall be made aware of defects and errors that may be encountered as part of their job functions. | Personnel Article 22 The manufacturer, etc. shall ensure that the personnel performing the work13 affecting the product quality are competent on the basis of the following requirements. (1) The appropriate training, and (2) The appropriate skills and experience. | No significant differences. |
| 6.2.2 Competence, awareness and training The organization shall a) Determine the necessary competence for personnel performing work affecting product quality, b) Provide training or take other actions to satisfy these needs, c) Evaluate the effectiveness of the actions taken, d) Ensure that its personnel are aware of the relevance and importance of their activities and how they contribute to the achievement of the quality objectives, and e) Maintain appropriate records of education, training, skills and experience (see 4.2.4). NOTE National or regional regulations might require the organization to establish documented procedures for identifying training needs. | 820.20 Management responsibility (b) Organization (2) Resources. Each manufacturer shall provide adequate resources, including the assignment of trained personnel, for management, performance of work, and assessment activities, including internal quality audits, to meet the requirements of this part. 820.25 Personnel. (a) General. Each manufacturer shall have sufficient personnel with the necessary education, background, training, and experience to assure that all activities required by this part are correctly performed. (b) Training. Each manufacturer shall establish procedures for identifying training needs and ensure that all personnel are trained to adequately perform their assigned responsibilities. Training shall be documented. (1) As part of their training, personnel shall be made aware of device defects, which may occur from the improper performance of their specific jobs. (2) Personnel who perform verification and validation activities shall be made aware of defects and errors that may be encountered as part of their job functions. | Competence, Awareness and Training Article 23 The manufacturer, etc. shall conduct the following duties. (1) To determine the necessary competence for the personnel performing the work affecting the product quality, (2) To establish the documented procedure for identifying the training needs, (3) To provide training or take other actions to satisfy the needs determined by the documented procedure specified in preceding Item (2), (4) To evaluate the effectiveness of the actions specified in preceding Item (3), (5) To ensure that the personnel are aware of the relevance and importance of their activities and how they contribute to the achievement of the quality objectives, and (6) To maintain appropriate records of the training, skills and experience. | No significant differences. |
| 6.3 Infrastructure The organization shall determine, provide and maintain the infrastructure needed to achieve conformity to product requirements. Infrastructure includes, as applicable a) Buildings, workspace and associated utilities, b) Process equipment (both hardware and software), and c) Supporting services (such as transport or communication). The organization shall establish documented requirements for maintenance activities, including their frequency, when such activities or lack thereof can affect product quality. Records of such maintenance shall be maintained (see 4.2.4)... | Production and process control (d) Personnel. Each manufacturer shall establish and maintain requirements for the health, cleanliness, personal practices, and clothing of personnel if contact between such personnel and product or environment could reasonably be expected to have an adverse effect on product quality. The manufacturer shall ensure that maintenance and other personnel who are required to work temporarily under special environmental conditions are appropriately trained or supervised by a trained individual. (f) Buildings. Buildings shall be of suitable design and contain sufficient space to perform necessary operations, prevent mix-ups, and assure orderly handling. (g) Equipment. Each manufacturer shall ensure that all equipment used in the manufacturing process meets specified requirements and is appropriately designed, constructed, placed, and installed to facilitate maintenance, adjustment, cleaning, and use. (1) Maintenance schedule. Each manufacturer shall establish and maintain schedules for the adjustment, cleaning, and other maintenance of equipment to ensure that manufacturing specifications are met. Maintenance activities, including the date and individual(s) performing the maintenance activities, shall be documented. (2) Inspection. Each manufacturer shall conduct periodic inspections in accordance with established procedures to ensure adherence to applicable equipment maintenance schedules. The inspections, including the date and individual(s) conducting the inspections, shall be documented. (3) Adjustment. Each manufacturer shall ensure that any inherent limitations or allowable tolerances are visibly posted on or near equipment requiring periodic adjustments or are readily available to personnel performing these adjustments. (h) Manufacturing material. Where a manufacturing material could reasonably be expected to have an adverse effect on product quality, the manufacturer shall establish and maintain procedures for the use and removal of such manufacturing material to ensure that it is removed or limited to an amount that does not adversely affect the device's quality. The removal or reduction of such manufacturing material shall be documented. | Infrastructure Article 24 1. The manufacturer, etc. shall determine, provide and maintain the following infrastructure, as applicable, needed to achieve the conformity to the product requirements. (1) The buildings, workspace and associated utilities, (2) The process equipment (both hardware and software), and (3) The supporting services (such as transport or communication). 2. The manufacturer, etc. shall provide and maintain the following infrastructure in case where they manufacture the following products. (1) For the products which need the control of dust, humidity, insects and rodents, the buildings or facilities for controlling dust, humidity, insects and rodents, (2) For the products of which manufacturing process the poisonous gases are handled, the facilities necessary for disposing of such poisonous gases, (3) For the products in a form of liquid, sol, gel or powder (excluding those concerned with the sterilised medical devices), the work rooms that meet the following requirements, 14 Being the buildings which do not allow passage of the personnel other than those conducting operations in such work rooms, with the proviso that this provision shall not apply in case where the products, etc. could not be contaminated by the personnel other than those conducting operations in such work rooms, b. Not being provided with the entrances directly leading to the outside (except those for emergency), with the proviso that this provision shall not apply in case where the work rooms are provided with the buildings and facilities necessary for preventing contamination due to the outside, Being provided with the entrances and windows that can be closed, Being provided with the buildings and facilities necessary for preventing contamination by dust or microorganisms according to the type and manufacturing process of the products, with the proviso that this provision shall not apply in case where the manufacturing facilities, etc. provide equivalent functions, Being provided the effluent facilities, in case where they are provided in the work rooms, with the buildings necessary for preventing contamination of the work rooms, and Being provided with the facilities for supplying water of the quality and quantity necessary for the manufacturing, according to the type and the manufacturing process of the products. 3. The manufacturer, etc. shall establish the documented requirements for the maintenance activities, including their frequency, when such activities or lack thereof can affect the product quality. 4. The manufacturer, etc. shall ensure that the records of such maintenance are maintained. | The intent of the three documents is consistent; the QSR contains a number of specific requirements related to the creation of maintenance schedules, inspections and adjustment of equipment, and manufacturing materials. The JGMP has more requirements relating to e.g. pest control, substance control, gasses etc. These requirements are logical and would, in my opinion, be expected by NB’s and FDA. |
| 6.4 Work environment The organization shall determine and manage the work environment needed to achieve conformity to product requirements. The following requirements shall apply. a) The organization shall establish documented requirements for health, cleanliness and clothing of personnel if contact between such personnel and the product or work environment could adversely affect the quality of the product (see 7.5.1.2.1). b) If work environment conditions can have an adverse effect on product quality, the organization shall establish documented requirements for the work environment conditions and documented procedures or work instructions to monitor and control these work environment conditions (see 7.5.1.2.1). c) The organization shall ensure that all personnel who are required to work temporarily under special environmental conditions within the work environment are appropriately trained or supervised by a trained person [see 6.2.2 b)]. d) If appropriate, special arrangements shall be established and documented for the control of contaminated or potentially contaminated product in order to prevent contamination of other product, the work environment or personnel (see 7.5.3.1). | Production and process controls (c) Environmental controls. Where environmental conditions could reasonably be expected to have an adverse effect on product quality, the manufacturer shall establish and maintain procedures to adequately control these environmental conditions. Environmental control system(s) shall be periodically inspected to verify that the system, including necessary equipment, is adequate and functioning properly. These activities shall be documented and reviewed (d) Personnel. Each manufacturer shall establish and maintain requirements for the health, cleanliness, personal practices, and clothing of personnel if contact between such personnel and product or environment could reasonably be expected to have an adverse effect on product quality. The manufacturer shall ensure that maintenance and other personnel who are required to work temporarily under special environmental conditions are appropriately trained or supervised by a trained individual. (e) Contamination control. Each manufacturer shall establish and maintain procedures to prevent contamination of equipment or product by substances that could reasonably be expected to have an adverse effect on product quality. | Work Environment Article 25 1. The manufacturer, etc. shall determine and manage the work environment needed to achieve the conformity to the product requirements. 2. The manufacturer, etc. shall establish the documented requirements for health, cleanliness and clothing of the personnel if contact between such personnel and the products, etc. or work environment could adversely affect the quality of the products, with the proviso that this provision shall not apply to the process in case where the products are cleaned in the subsequent process in accordance with the provision of Item (1) or (2) of Paragraph 1 of Article 41. 3. If the work environment conditions can have an adverse effect on the product quality, the manufacturer, etc. shall establish the documented requirements for the work environment conditions and the documented procedure or work instructions to monitor and control these work environment conditions, with the proviso that this provision15 shall not apply to the process in case where the products are cleaned in the subsequent process in accordance with the provision of Item (1) or (2) of Paragraph 1 of Article 41. 4. The manufacturer, etc. shall ensure that all personnel who are required to work temporarily under special environmental conditions within the work environment are appropriately trained as specified in Item (3) of Article 23, with the proviso that this provision shall not apply in case where the personnel are supervised by a trained person. 5. The manufacturer, etc. shall ensure, if appropriate, that the special arrangements are established and documented for the control (including identification specified in Paragraph 3 of Article 47) of the contaminated or potentially contaminated products in order to prevent the contamination of other products, the work environment or personnel. | The intent of the three documents is consistent; ISO/DIS 13485 and the JGMP specifically calls for control of used product to prevent contamination of other product, the manufacturing environment or personnel. |
| 7 Product realization 7.1 Planning of product realization The organization shall plan and develop the processes needed for product realization. Planning of product realization shall be consistent with the requirements of the other processes of the quality management system (see 4.1). In planning product realization, the organization shall determine the following, as appropriate: a) Quality objectives and requirements for the product; b) The need to establish processes, documents, and provide resources specific to the product; c) Required verification, validation, monitoring, inspection and test activities specific to the product and the criteria for product acceptance; d) Records needed to provide evidence that the realization processes and resulting product meet requirements (see 4.2.4). The output of this planning shall be in a form suitable for the organization’s method of operations. The organization shall establish documented requirements for risk management throughout product realization. Records arising from risk management shall be maintained (see 4.2.4). NOTE 1 A document specifying the processes of the quality management system (including the product realization processes) and the resources to be applied to a specific product, project or contract, can be referred to as a quality plan. NOTE 2 The organization may also apply the requirements given in 7.3 to the development of product realization Processes. NOTE 3 See ISO 14971 for guidance related to risk management. | 820.5 Quality system. Each manufacturer shall establish and maintain a quality system that is appropriate for the specific medical device(s) designed or manufactured, and that meets the requirements of this part. | Section 5 Product Realisation Planning of Product Realisation Article 26 1. The manufacturer, etc. shall plan and develop the processes needed for the product realisation (the duties of the manufacturer, etc. for realisation of the products). 2. The manufacturer, etc. shall ensure that the planning of the product realisation specified in preceding Paragraph 1 (hereinafter referred to as “planning of product realisation”) is consistent with the requirements of the other processes of the quality management system. The manufacturer, etc. in case where they conduct planning of product realisation, shall determine the following items, as appropriate. The quality objectives and requirements for the products, The need to establish the processes, QMS documents, and provide resources specific to the products, The required verification, validation, monitoring, inspection and test activities specific to the products and the criteria for the product acceptance (hereinafter referred to as the “manufacturing site release criteria”), The records needed to provide evidence that the realisation processes and resulting products meet the product requirements. 4. The manufacturer, etc. shall ensure that the output of the planning of product realisation is in a form suitable for the manufacturing site's method of the operations. 5. The manufacturer, etc. shall establish the documented requirements for the risk management throughout the product realisation. 6. The manufacturer, etc. shall ensure that the records arising from the risk management are maintained.16 | The QSR doesn’t specifically recognize the concept of product realization, even though it includes requirements related to essentially all of the processes associated with product realization. The QSR also requires, either implicitly or explicitly, in various sections planning associated with these requirements. ISO/DIS 13485:2003 recognizes the fact that risk management is a process that shall be conducted throughout the product realization process, while the QSR refers to risk management only in the section related to design validation. FDA, however, recognizes the wisdom of the ISO/DIS 13485:2003 risk management requirements will seek records associated with risk management consistent with the requirements set out in ISO/DIS 13485:2003. |
| 7.2 Customer-related processes 7.2.1 Determination of requirements related to the product The organization shall determine a) Requirements specified by the customer, including the requirements for delivery and post-delivery activities, b) Requirements not stated by the customer but necessary for specified or intended use, where known, c) Statutory and regulatory requirements related to the product, and d) Any additional requirements determined by the organization. |  | Determination of Requirements Related to the Products Article 27 The manufacturer, etc. shall determine the following items as the product requirements. (1) The requirements specified by the customer, including the requirements for delivery and post-delivery activities, (2) The requirements not stated by the customer but necessary for the specified or intended use, where known, (3) The regulatory requirements related to the products, and (4) Any additional requirements determined by the manufacturer, etc. | The closest the QSR gets to determine customer requirements related to the product in 820.30(c) design inputs. The customer requirements referred to in clause 7.2 of ISO/DIS 13485:2003 and article 27 of the JGMP refer to those requirements associated with getting the product to the customer. This includes items associated with order handling and what were referred to in early versions of ISO 9001 as contract review. These are not focuses of the QS Reg. |
| 7.2.2 Review of requirements related to the product The organization shall review the requirements related to the product. This review shall be conducted prior to the organization's commitment to supply a product to the customer (e.g. submission of tenders, acceptance of contracts or orders, acceptance of changes to contracts or orders) and shall ensure that a) Product requirements are defined and documented, b) Contract or order requirements differing from those previously expressed are resolved, and c) The organization has the ability to meet the defined requirements. Records of the results of the review and actions arising from the review shall be maintained (see 4.2.4). Where the customer provides no documented statement of requirement, the customer requirements shall be confirmed by the organization before acceptance. Where product requirements are changed, the organization shall ensure that relevant documents are amended and that relevant personnel are made aware of the changed requirements. NOTE In some situations, such as internet sales, a formal review is impractical for each order. Instead the review can cover relevant product information such as catalogues or advertising material. |  | Review of Requirements Related to Products Article 28 1. The manufacturer, etc. shall review the product requirements, prior to their commitment to supply the products to the customer. 2. The manufacturer, etc. shall, in case where they conduct the review specified in preceding Paragraph 1, ensure the following items. (1) To be ensured that the product requirements are defined and documented, (2) To be ensured that the contract or order requirements of the customer differing from those previously expressed are resolved, and (3) To be ensured that the manufacturing site has the ability to meet the defined requirements. 3. The manufacturer, etc. shall ensure that the records of the results of the review specified in preceding Paragraph 1 and the actions arising from the review are maintained. 4. Where the customer provides no documented statement of the requirement, the manufacturer, etc. shall confirm the customer requirements before the acceptance. 5. Where the product requirements are changed, the manufacturer, etc. shall ensure that the relevant documents are amended and that relevant personnel are made aware of the changed requirements. |  |
| 7.2.3 Customer communication The organization shall determine and implement effective arrangements for communicating with customers in relation to a) Product information, b) Enquiries, contracts or order handling, including amendments, c) Customer feedback, including customer complaints (see 8.2.1), and d) Advisory notices (see 8.5.1). |  | Customer Communication Article 29 The manufacturer, etc. shall determine and implement the effective arrangements for communicating with the customers in relation to the following items. (1) The product information, (2) The enquiries, contracts or order handling, including the amendments, (3) The customer feedback, including the customer complaints, and17 (4) The issuance and implementation of the advisory notices specified in Paragraph 2 of Article 62. |  |
| 7.3 Design and development 7.3.1 Design and development planning The organization shall establish documented procedures for design and development. The organization shall plan and control the design and development of product. During the design and development planning, the organization shall determine a) The design and development stages, b) The review, verification, validation and design transfer activities (see Note) that are appropriate at each design and development stage, and c) The responsibilities and authorities for design and development. The organization shall manage the interfaces between different groups involved in design and development to ensure effective communication and clear assignment of responsibility. Planning output shall be documented, and updated as appropriate, as the design and development progresses (see 4.2.3). NOTE Design transfer activities during the design and development process ensure that design and development outputs are verified as suitable for manufacturing before becoming final production specifications. | Design controls (a) General. (1) Each manufacturer of any class III or class II device, and the class I devices listed in paragraph (a)(2) of this section, shall establish and maintain procedures to control the design of the device in order to ensure that specified design requirements are met. (2) The following class I devices are subject to design controls: (i) Devices automated with computer software; and (ii) The devices listed in the following chart. -------------------------------------------------------- Section Device -------------------------------------------------------- 868.6810...Catheter, Tracheobronchial Suction. 878.4460... Glove, Surgeon's. 880.6760... Restraint, Protective. 892.5650... System, Applicator, Radionuclide, Manual. 892.5740... Source, Radionuclide Teletherapy. -------------------------------------------------------- (b) Design and development planning. Each manufacturer shall establish and maintain plans that describe or reference the design and development activities and define responsibility for implementation. The plans shall identify and describe the interfaces with different groups or activities that provide, or result in, input to the design and development process. The plans shall be reviewed, updated, and approved as design and development evolves. (j) Design history file. Each manufacturer shall establish and maintain a DHF for each type of device. The DHF shall contain or reference the records necessary to demonstrate that the design was developed in accordance with the approved design plan and the requirements of this part. | Design and Development Planning Article 30 1. The manufacturer, etc. shall establish the documented procedure for the design and development. 2. The manufacturer, etc. shall establish the plan of the design and development of the products (hereinafter referred to as the “design and development plan”) and control the design and development of the products. 3. The manufacturer, etc. when they establish the design and development plan, shall determine the following items. (1) The design and development stages, (2) The review, verification, validation and design transfer activities (the duties, during the design and development process, to ensure that the design and development outputs are verified as suitable for the manufacturing before becoming the final production specifications) that are appropriate at each design and development stage, and (3) The responsibilities and authorities of the departments and personnel concerned with the design and development. 4. The manufacturer, etc. shall manage the interfaces between different groups involved in the design and development to ensure effective communication and clear assignment of the responsibility. 5. The manufacturer, etc. shall ensure that the design and development plan are documented, and updated as appropriate, as the design and development progresses. | The overall objectives of the three documents related to design control planning are consistent. The QSR limits the applicability of design controls to more high-risk medical devices, while ISO/DIS 13485:2003 and JGMP applies them to all medical devices. While ISO/DIS 13485:2003 and JGMP requires the creation of design control documentation and records, the QSR requires the establishment of a Design History File, which either contains or refers to all the documents and records associated with the application of the design control processes to a particular product. |
| 7.3.2 Design and development inputs Inputs relating to product requirements shall be determined and records maintained (see 4.2.4). These inputs shall include a) Functional, performance and safety requirements, according to the intended use, b) Applicable statutory and regulatory requirements, c) Where applicable, information derived from previous similar designs, d) Other requirements essential for design and development, and e) Output(s) of risk management (see 7.1). These inputs shall be reviewed for adequacy and approved. Requirements shall be complete, unambiguous and not in conflict with each other. | Design controls (c) Design input. Each manufacturer shall establish and maintain procedures to ensure that the design requirements relating to a device are appropriate and address the intended use of the device, including the needs of the user and patient. The procedures shall include a mechanism for addressing incomplete, ambiguous, or conflicting requirements. The design input requirements shall be documented and shall be reviewed and approved by a designated individual(s). The approval, including the date and signature of the individual(s) approving the requirements, shall be documented. | Design and Development Inputs Article 31 1. The manufacturer, etc. shall ensure that the following inputs relating to the product requirements are determined and the records are maintained. (1) The functional, performance and safety requirements, according to the intended use, (2) Where applicable, the information derived from the previous similar designs, (3) The outputs of the risk management specified in Paragraph 5 of Article 26, (4) The applicable regulatory requirements, and (5) Other requirements essential for the design and development.18 2. The manufacturer, etc. shall ensure that these inputs are reviewed for adequacy and approved. | No significant differences. |
| 7.3.3 Design and development outputs The outputs of design and development shall be provided in a form that enables verification against the design and development input and shall be approved prior to release. Design and development outputs shall a) Meet the input requirements for design and development, b) Provide appropriate information for purchasing, production and for service provision, c) Contain or reference product acceptance criteria, and d) Specify the characteristics of the product that are essential for its safe and proper use. Records of the design and development outputs shall be maintained (see 4.2.4). NOTE Records of design and development outputs can include specifications, manufacturing procedures, engineering drawings, and engineering or research logbooks. | Design controls (d) Design output. Each manufacturer shall establish and maintain procedures for defining and documenting design output in terms that allow an adequate evaluation of conformance to design input requirements. Design output procedures shall contain or make reference to acceptance criteria and shall ensure that those design outputs that are essential for the proper functioning of the device are identified. Design output shall be documented, reviewed, and approved before release. The approval, including the date and signature of the individual(s) approving the output, shall be documented. | Design and Development Outputs Article 32 1. The manufacturer, etc. shall ensure that the outputs of the design and development are provided in a form that enables the verification against the design and development input. 2. The manufacturer, etc. shall ensure that the outputs of the design and development are approved prior to the release to the next process. 3. The manufacturer, etc. shall ensure that the design and development outputs meet the following requirements. (1) To meet the input requirements for the design and development, (2) To provide appropriate information for the purchasing, production and for the service provision, (3) To contain or reference the product acceptance criteria, and (4) To specify the characteristics of the products that are essential for their safe and proper use. 4. The manufacturer, etc. shall ensure that the records of the design and development outputs are maintained. | The intent of the three documents is consistent, with the QSR containing specific requirements associated with the approval and release of design outputs. |
| 7.3.4 Design and development review At suitable stages, systematic reviews of design and development shall be performed in accordance with planned arrangements (see 7.3.1) a) To evaluate the ability of the results of design and development to meet requirements, and b) To identify any problems and propose necessary actions. Participants in such reviews shall include representatives of functions concerned with the design and development stage(s) being reviewed, as well as other specialist personnel (see 5.5.1 and 6.2.1). Records of the results of the reviews and any necessary actions shall be maintained (see 4.2.4). | Design controls (e) Design review. Each manufacturer shall establish and maintain procedures to ensure that formal documented reviews of the design results are planned and conducted at appropriate stages of the device's design development. The procedures shall ensure that participants at each design review include representatives of all functions concerned with the design stage being reviewed and an individual(s) who does not have direct responsibility for the design stage being reviewed, as well as any specialists needed. The results of a design review, including identification of the design, the date, and the individual(s) performing the review, shall be documented in the design history file (the DHF). | Design and Development Review Article 33 1. The manufacturer, etc. shall ensure that, at suitable stages, the systematic reviews of the design and development (hereinafter referred to as “design and development review”) are performed for the purpose of the following items in accordance with the arrangements defined in the design and development plan. (1) To evaluate the ability of the results of the design and development to meet the requirements, and (2) To identify any problems and propose necessary actions. 2. The manufacturer, etc. shall ensure that the participants in such reviews include the representatives of the departments concerned with the design and development stages being reviewed, as well as other specialist personnel. 3. The manufacturer, etc. shall ensure that the records of the results of the reviews and any necessary actions are maintained. | The intent of the three documents is consistent, with ISO/DIS 13485:2003 and JGMP illustrating in more detail the objectives of the design review and the QSR including prescriptive design review process requirements not contained in the standard. |
| 7.3.5 Design and development verification Verification shall be performed in accordance with planned arrangements (see 7.3.1) to ensure that the design and development outputs have met the design and development input requirements. Records of the results of the verification and any necessary actions shall be maintained (see 4.2.4). | Design controls (f) Design verification. Each manufacturer shall establish and maintain procedures for verifying the device design. Design verification shall confirm that the design output meets the design input requirements. The results of the design verification, including identification of the design, method(s), the date, and the individual(s) performing the verification, shall be documented in the DHF. | Design and Development Verification19 Article 34 1. The manufacturer, etc. shall ensure that the verification is performed in accordance with the arrangements defined in the design and development plan to ensure that the design and development outputs have met the design and development input requirements. 2. The manufacturer, etc. shall ensure that the records of the results of the verification specified in preceding Paragraph 1 and any necessary actions taken are maintained. | The intent of the three documents is consistent, with the QSR including prescriptive design verification process requirements not contained in 13485 and JGMP. |
| 7.3.6 Design and development validation Design and development validation shall be performed in accordance with planned arrangements (see 7.3.1) to ensure that the resulting product is capable of meeting the requirements for the specified application or intended use. Validation shall be completed prior to the delivery or implementation of the product (see Note 1). Records of the results of validation and any necessary actions shall be maintained (see 4.2.4). As part of design and development validation, the organization shall perform clinical evaluations and/or evaluation of performance of the medical device, as required by national or regional regulations (see Note 2). NOTE 1 If a medical device can only be validated following assembly and installation at point of use, delivery is not considered to be complete until the product has been formally transferred to the customer. NOTE 2 Provision of the medical device for purposes of clinical evaluations and/or evaluation of performance is not considered to be delivery. | Design controls (g) Design validation. Each manufacturer shall establish and maintain procedures for validating the device design. Design validation shall be performed under defined operating conditions on initial production units, lots, or batches, or their equivalents. Design validation shall ensure that devices conform to defined user needs and intended uses and shall include testing of production units under actual or simulated use conditions. Design validation shall include software validation and risk analysis, where appropriate. The results of the design validation, including identification of the design, method(s), the date, and the individual(s) performing the validation, shall be documented in the DHF. | Design and Development Validation Article 35 1. The manufacturer, etc. shall ensure that the design and development validation (hereinafter referred to as “design and development validation”) is performed in accordance with the arrangements defined in the design and development plan to ensure that the resulting products are capable of meeting the requirements for the specified application or intended use. 2. The manufacturer, etc. shall ensure that the design and development validation is completed prior to the delivery or implementation of the products. In case where the design and development validation can only be conducted following the assembly and installation at the point of the use, the validation shall be completed prior to the delivery of the medical devices to their user. 3. The manufacturer, etc. shall ensure that the records of the results of the design and development validation and any necessary actions are maintained. | The intent of the three documents is consistent. The QSR seems to indicate that risk analysis is a design validation process. This is not consistent with the teachings of ISO 14971:2000, which calls for risk management activities throughout the product realization process. ISO/DIS 13485:3002 and JGMP addresses a scenario not addressed by the QSR (e.g., where final assembly of the medical device is accomplished on delivery to the customer) The QSR contains a number of design validation process requirements not included in the standard. |
| 7.3.7 Control of design and development changes Design and development changes shall be identified and records maintained. The changes shall be reviewed, verified and validated, as appropriate, and approved before implementation. The review of design and development changes shall include evaluation of the effect of the changes on constituent parts and product already delivered. Records of the results of the review of changes and any necessary actions shall be maintained (see 4.2.4). | Design controls (i) Design changes. Each manufacturer shall establish and maintain procedures for the identification, documentation, and validation or where appropriate verification, review, and approval of design changes before their implementation. | Control of Design and Development Changes Article 36 1. The manufacturer, etc. shall ensure that the design and development changes are identified and records thereof are maintained. 2. The manufacturer, etc. shall ensure that the changes are reviewed, verified and validated, as appropriate, and approved before the implementation. 3. The manufacturer, etc. shall ensure that the review of the design and development changes includes the evaluation of the effect of the changes on the constituent parts and products already delivered. 4. The manufacturer, etc. shall ensure that the records of the results of the review of the changes specified in preceding Paragraph 2 and any necessary actions are maintained. | ISO/DIS 13485:2003 contains requirements related to the effect of design changes on product already delivered and records of design changes that do not appear in the QSR. |
| 7.4 Purchasing 7.4.1 Purchasing process The organization shall establish documented procedures to ensure that purchased product conforms to specified purchase requirements. The type and extent of control applied to the supplier and the purchased product shall be dependent upon the effect of the purchased product on subsequent product realization or the final product. The organization shall evaluate and select suppliers based on their ability to supply product in accordance with the organization’s requirements. Criteria for selection, evaluation and re-evaluation shall be established. Records of the results of evaluations and any necessary actions arising from the evaluation shall be maintained (see 4.2.4). | 820.50 Purchasing controls. Each manufacturer shall establish and maintain procedures to ensure that all purchased or otherwise received product and services conform to specified requirements. (a) Evaluation of suppliers, contractors, and consultants. Each manufacturer shall establish and maintain the requirements, including quality requirements that must be met by suppliers, contractors, and consultants. Each manufacturer shall: (1) Evaluate and select potential suppliers, contractors, and consultants on the basis of their ability to meet specified requirements, including quality requirements. The evaluation shall be documented. (2) Define the type and extent of control to be exercised over the product, services, suppliers, contractors, and consultants, based on the evaluation results. (3) Establish and maintain records of acceptable suppliers, contractors, and consultants. | Purchasing Process Article 37 1. The manufacturer, etc. shall establish the documented procedure to ensure that the purchased products conform to the requirements that they specify for the purchased products (hereinafter referred to as “purchased product requirements”). 2. The manufacturer, etc. shall ensure that the type and extent of the control applied to the supplier and the purchased products are dependent upon the effect of the purchased products on subsequent product realisation or the final products.20 3. The manufacturer, etc. shall evaluate and select the suppliers based on their ability to supply the products in accordance with the purchased product requirements. 4. The manufacturer, etc. shall ensure that the criteria for the selection, evaluation and re-evaluation are established. 5. The manufacturer, etc. shall ensure that the records of the results of the evaluations specified in preceding Paragraph 3 and any necessary actions arising from the evaluation are maintained. | No significant differences. |
| 7.4.2 Purchasing information Purchasing information shall describe the product to be purchased, including where appropriate a) Requirements for approval of product, procedures, processes and equipment, b) Requirements for qualification of personnel, and c) Quality management system requirements. The organization shall ensure the adequacy of specified purchase requirements prior to their communication to the supplier. To the extent required for traceability given in 7.5.3.2, the organization shall maintain relevant purchasing information, i.e. documents (see 4.2.3) and records (see 4.2.4). | Purchasing controls (b) Purchasing data. Each manufacturer shall establish and maintain data that clearly describe or reference the specified requirements, including quality requirements, for purchased or otherwise received product and services. Purchasing documents shall include, where possible, an agreement that the suppliers, contractors, and consultants agree to notify the manufacturer of changes in the product or service so that manufacturers may determine whether the changes may affect the quality of a finished device. Purchasing data shall be approved in accordance with Sec. 820.40. | Purchasing Information Article 38 1. The manufacturer, etc. shall ensure that the information on the purchased products (hereinafter referred to as “purchasing information”) describes the products to be purchased, including the following items where appropriate. (1) The requirements for the approval of the purchased products, procedures, processes and equipment, (2) The requirements for the qualification of the personnel, (3) The quality management system requirements, and (4) Other information necessary for the purchased products. 2. The manufacturer, etc. shall ensure the adequacy of the purchased product requirements prior to their communication to the supplier. 3. The manufacturer, etc. shall maintain, to the extent required for the traceability given in the documented procedure in accordance with the provision of Paragraph 2 of Article 48, relevant purchasing information, i.e. documents and records. | The intent of all three documents is consistent. ISO/DIS 13485:2003 and the JGMP contain a requirement associated with the organization ensuring the adequacy of purchasing requirements prior to communicating them to the supplier. The QSR contains a requirement that the organization obtain, where possible, the agreement of the supplier to notify the organization of changes to the product or service so that the organization can assess the potential effect on the quality of the medical device. The JGMP specifies that the manufacturer shall obtain other necessary information regarding purchased products, thereby ensuring that one does not claim ignorance to things that one should have investigated. |
| 7.4.3 Verification of purchased product The organization shall establish and implement the inspection or other activities necessary for ensuring that purchased product meets specified purchase requirements. Where the organization or its customer intends to perform verification at the supplier’s premises, the organization shall state the intended verification arrangements and method of product release in the purchasing information. Records of the verification shall be maintained (see 4.2.4). | 820.80 Receiving, in-process, and finished device acceptance (b) Receiving acceptance activities. Each manufacturer shall establish and maintain procedures for acceptance of incoming product. Incoming product shall be inspected, tested, or otherwise verified as conforming to specified requirements. Acceptance or rejection shall be documented. | Verification of Purchased Products Article 39 1. The manufacturer, etc. shall establish and implement the inspection or other activities necessary for ensuring that the purchased products meet the specified purchase requirements. 2. Where the manufacturer, etc. or their customer intends to perform the verification at the supplier's premises, the manufacturer, etc. shall state the intended verification arrangements and method of the product release in the purchasing information specified in preceding Article. 3. The manufacturer, etc. shall maintain the records of the verification specified in preceding Paragraph 2. | ISO/DIS 13485:2003 and the JGMP contain requirements related to the scenario where the organization seeks to verify purchased product at the supplier’s location. |
| 7.5 Production and service provision 7.5.1 Control of production and service provision 7.5.1.1 General requirements The organization shall plan and carry out production and service provision under controlled conditions. Controlled conditions shall include, as applicable a) The availability of information that describes the characteristics of the product, b) The availability of documented procedures, documented requirements, work instructions, and reference materials and reference measurement procedures as necessary, c) The use of suitable equipment, d) The availability and use of monitoring and measuring devices, e) The implementation of monitoring and measurement, f) The implementation of release, delivery and post-delivery activities, and g) The implementation of defined operations for labelling and packaging. The organization shall establish and maintain a record (see 4.2.4) for each batch of medical devices that provides traceability to the extent specified in 7.5.3 and identifies the amount manufactured and amount approved for distribution. The batch record shall be verified and approved. NOTE A batch can be a single medical device. 7.5.1.2 Control of production and service provision — Specific requirements 7.5.1.2.1 Cleanliness of product and contamination control The organization shall establish documented requirements for cleanliness of product if a) Product is cleaned by the organization prior to sterilization and/or its use, or b) Product is supplied non-sterile to be subjected to a cleaning process prior to sterilization and/or its use, or c) Product is supplied to be used non-sterile and its cleanliness is of significance in use, or d) Process agents are to be removed from product during manufacture. If product is cleaned in accordance with a) or b) above, the requirements contained in 6.4 a) and 6.4 b) do not apply prior to the cleaning process. 7.5.1.2.2 Installation activities If appropriate, the organization shall establish documented requirements, which contain acceptance criteria for installing and verifying the installation of the medical device. If the agreed customer requirements allow installation to be performed other than by the organization or its authorized agent, the organization shall provide documented requirements for installation and verification. Records of installation and verification performed by the organization or its authorized agent shall be maintained (see 4.2.4). 7.5.1.2.3 Servicing activities If servicing is a specified requirement, the organization shall establish documented procedures, work instructions and reference materials and reference measurement procedures, as necessary, for performing servicing activities and verifying that they meet the specified requirements. Records of servicing activities carried out by the organization shall be maintained (see 4.2.4). NOTE Servicing can include, for example, repair and maintenance. 7.5.1.3 Particular requirements for sterile medical devices The organization shall maintain records of the process parameters for the sterilization process, which was used for each sterilization batch (see 4.2.4). Sterilization records shall be traceable to each production batch of medical devices (see 7.5.1.1). | 820.70 Production and process controls. General. Each manufacturer shall develop, conduct, control, and monitor production processes to ensure that a device conforms to its specifications. Where deviations from device specifications could occur as a result of the manufacturing process, the manufacturer shall establish and maintain process control procedures that describe any process controls necessary to ensure conformance to specifications. Where process controls are needed they shall include: 1) Documented instructions, standard operating procedures (SOP's), and methods that define and control the manner of production; (2) Monitoring and control of process parameters and component and device characteristics during production; (3) Compliance with specified reference standards or codes; (4) The approval of processes and process equipment; and (5) Criteria for workmanship, which shall be expressed in documented standards or by means of, identified and approved representative samples. (b) Production and process changes. Each manufacturer shall establish and maintain procedures for changes to a specification, method, process, or procedure. Such changes shall be verified or where appropriate validated according to Sec. 820.75, before implementation and these activities shall be documented. Changes shall be approved in accordance with Sec. 820.40. (c) Environmental controls. Where environmental conditions could reasonably be expected to have an adverse effect on product quality, the manufacturer shall establish and maintain procedures to adequately control these environmental conditions. Environmental control system(s) shall be periodically inspected to verify that the system, including necessary equipment, is adequate and functioning properly. These activities shall be documented and reviewed. (d) Personnel. Each manufacturer shall establish and maintain requirements for the health, cleanliness, personal practices, and clothing of personnel if contact between such personnel and product or environment could reasonably be expected to have an adverse effect on product quality. The manufacturer shall ensure that maintenance and other personnel who are required to work temporarily under special environmental conditions are appropriately trained or supervised by a trained individual. (e) Contamination control. Each manufacturer shall establish and maintain procedures to prevent contamination of equipment or product by substances that could reasonably be expected to have an adverse effect on product quality. (f) Buildings. Buildings shall be of suitable design and contain sufficient space to perform necessary operations, prevent mix-ups, and assure orderly handling. (g) Equipment. Each manufacturer shall ensure that all equipment used in the manufacturing process meets specified requirements and is appropriately designed, constructed, placed, and installed to facilitate maintenance, adjustment, cleaning, and use. (1) Maintenance schedule. Each manufacturer shall establish and maintain schedules for the adjustment, cleaning, and other maintenance of equipment to ensure that manufacturing specifications are met. Maintenance activities, including the date and individual(s) performing the maintenance activities, shall be documented. (2) Inspection. Each manufacturer shall conduct periodic inspections in accordance with established procedures to ensure adherence to applicable equipment maintenance schedules. The inspections, including the date and individual(s) conducting the inspections, shall be documented. (3) Adjustment. Each manufacturer shall ensure that any inherent limitations or allowable tolerances are visibly posted on or near equipment requiring periodic adjustments or are readily available to personnel performing these adjustments. (h) Manufacturing material. Where a manufacturing material could reasonably be expected to have an adverse effect on product quality, the manufacturer shall establish and maintain procedures for the use and removal of such manufacturing material to ensure that it is removed or limited to an amount that does not adversely affect the device's quality. The removal or reduction of such manufacturing material shall be documented. (i) Automated processes. When computers or automated data processing systems are used as part of production or the quality system, the manufacturer shall validate computer software for its intended use according to an established protocol. All software changes shall be validated before approval and issuance. These validation activities and results shall be documented. 820.170 Installation. (a) Each manufacturer of a device requiring installation shall establish and maintain adequate installation and inspection instructions, and where appropriate test procedures. Instructions and procedures shall include directions for ensuring proper installation so that the device will perform as intended after installation. The manufacturer shall distribute the instructions and procedures with the device or otherwise make them available to the person(s) installing the device. (b) The person installing the device shall ensure that the installation, inspection, and any required testing are performed in accordance with the manufacturer's instructions and procedures and shall document the inspection and any test results to demonstrate proper installation. 820.200 Servicing. (a) Where servicing is a specified requirement, each manufacturer shall establish and maintain instructions and procedures for performing and verifying that the servicing meets the specified requirements. (b) Each manufacturer shall analyze service reports with appropriate statistical methodology in accordance with Sec. 820.100. (c) Each manufacturer who receives a service report that represents an event, which must be reported to FDA under part 803 or 804 of this chapter, shall automatically consider the report a complaint and shall process it in accordance with the requirements of Sec. 820.198. (d) Service reports shall be documented and shall include: (1) The name of the device serviced; (2) Any device identification(s) and control number(s) used; (3) The date of service; (4) The individual(s) servicing the device; (5) The service performed; and (6) The test and inspection data. Design controls (h) Design transfer. Each manufacturer shall establish and maintain procedures to ensure that the device design is correctly translated into production specifications. | Control of Production and Service Provision Article 40 1. The manufacturer, etc. shall plan and carry out the production and service21 provision under the controlled conditions including the following, as applicable. (1) The availability of the information that describes the characteristics of the products, (2) The availability of the documented procedure, documented requirements, work instructions, and reference materials and reference measurement procedures as necessary, (3) The use of the suitable equipment, (4) The availability and use of the monitoring and measuring devices, (5) The implementation of the monitoring and measurement in accordance with the provisions of Articles 57 to 59, (6) The implementation of the release, delivery and post-delivery activities, and (7) The implementation of the defined operations for labelling and packaging defined in the documented procedure, etc. 2. The manufacturer, etc. shall establish and maintain records for each batch of the products that provides traceability to the extent defined in the documented procedure in accordance with the provision of Paragraph 2 of Article 48 and identifies the amount manufactured and amount approved for distribution. The batch record established in accordance with the provision of preceding Paragraph 2 shall be verified and approved. Cleanliness of Products and Contamination Control Article 41 The manufacturer, etc. shall establish the documented requirements for the cleanliness of the products in case where they correspond to the following cases. (1) The products are cleaned by the manufacturer, etc. prior to the sterilisation and/or their use, (2) The products are supplied non-sterile to be subjected to the cleaning process prior to the sterilisation and/or their use, (3) The products are supplied to be used non-sterile and their cleanliness is of significance in use, or (4) The process agents are to be removed from the products during the manufacture. Installation Activities Article 42 1. If appropriate, the manufacturer, etc. shall, in case where they manufacture the products concerned with the installation-control-required medical devices specified in Paragraph 1 of Article 93 of Enforcement Regulations, establish the documented22 requirements which contain the acceptance criteria for installing and verifying the installation of the medical devices. 2. The manufacturer, etc. shall provide the documented requirements for the installation and verification specified in preceding Paragraph 1 to the marketing authorisation holder. Servicing Activities Article 43 If the servicing is the specified requirement, the manufacturer, etc. shall establish the documented procedure, work instructions and reference materials and reference measurement procedures, as necessary, for performing the servicing activities (hereinafter referred to as “servicing activities”) and verifying that they meet the specified requirements. The manufacturer, etc. shall ensure that the records of the servicing activities carried out by them are maintained. | The intent of all three documents is consistent with each of the documents providing details of control or the types of processes that must be controlled in a way that supplements each other. It is suggested that the sections be read together in order to get a complete list of processes and process controls that are to be included in the quality management system Here is no specific reference to design transfer in ISO/DIS 13485:2003. |
| 7.5.2 Validation of processes for production and service provision 7.5.2.1 General requirements The organization shall validate any processes for production and service provision where the resulting output cannot be verified by subsequent monitoring or measurement. This includes any processes where deficiencies become apparent only after the product is in use or the service has been delivered. Validation shall demonstrate the ability of these processes to achieve planned results. The organization shall establish arrangements for these processes including, as applicable a) Defined criteria for review and approval of the processes, b) Approval of equipment and qualification of personnel, c) Use of specific methods and procedures, d) Requirements for records (see 4.2.4), and e) Revalidation. The organization shall establish documented procedures for the validation of the application of computer software (and changes to such software and/or its application) for production and service provision that affect the ability of the product to conform to specified requirements. Such software applications shall be validated prior to initial use. Records of validation shall be maintained (see 4.2.4) 7.5.2.2 Particular requirements for sterile medical devices The organization shall establish documented procedures for the validation of sterilization processes. Sterilization processes shall be validated prior to initial use. Records of validation of each sterilization process shall be maintained (see 4.2.4). | 820.75 Process validation. (a) Where the results of a process cannot be fully verified by subsequent inspection and test, the process shall be validated with a high degree of assurance and approved according to established procedures. The validation activities and results, including the date and signature of the individual(s) approving the validation and where appropriate the major equipment validated, shall be documented. (b) Each manufacturer shall establish and maintain procedures for monitoring and control of process parameters for validated processes to ensure that the specified requirements continue to be met. (1) Each manufacturer shall ensure that validated processes are performed by qualified individual(s). (2) For validated processes, the monitoring and control methods and data, the date performed, and, where appropriate, the individual(s) performing the process or the major equipment used shall be documented. (c) When changes or process deviations occur, the manufacturer shall review and evaluate the process and perform revalidation where appropriate. These activities shall be documented. | Validation of Processes for Production and Service Provision Article 45 The manufacturer, etc. shall validate any process for the production and service provision where the resulting output cannot be verified by the subsequent monitoring or measurement. This includes any processes where deficiencies become apparent only after the products are in use or the service has been delivered. The manufacturer, etc. shall ensure that the validation demonstrates the ability of the processes specified in preceding Paragraph 1 to achieve the planned results specified in Paragraph 1 of Article 14. The manufacturer, etc. shall establish the arrangements for the processes which should be validated in accordance with the provision of preceding Paragraph 1 including the following items, as applicable The defined criteria for the review and approval of the processes, The approval of the equipment and the qualification of the personnel, The use of the specific methods and procedures, The requirements for the records specified in Article 9, and The revalidation (conducting the validation again in case where the manufacturing procedure has been changed, etc.). The manufacturer, etc. shall establish the documented procedure for the validation of the application of computer software (and changes to such software and/or its application) for the production and service provision that affect the ability of the products to conform to the specified requirements. The manufacturer, etc. shall ensure that such software applications are validated prior to the initial use. The manufacturer, etc. shall ensure that the records of the validation specified in preceding Paragraphs 1 to 5 are maintained. Manufacturing Control of Sterile Medical Devices Article 44 The manufacturer in the category specified in Item (2) of Paragraph 5 of Article 26 of Enforcement Regulations or the foreign manufacturer in the category specified in Item (2) of Paragraph 4 of Article 36 of Enforcement Regulations (hereinafter collectively referred to as “sterile-medical-device-category manufacturer, etc.”) shall maintain the records of the process parameters for the sterilisation process which was used for each sterilisation lot. The sterile-medical-device-category manufacturer, etc. shall ensure that the sterilisation records specified in preceding Paragraph 1 are traceable to each production lot of the products. The sterile-medical-device-category manufacturer, etc. shall provide and maintain the following infrastructure in addition to those specified in Paragraphs 1 and 2 of Article 24. The buildings and facilities necessary for preventing contamination with dust or micro organisms, according to the manufacturing process of the products, with the proviso that this provision shall not apply in case where the manufacturing facilities, etc. provide equivalent functions, The buildings and facilities for maintaining the degree of cleanliness according to the manufacturing process of the products in the work rooms or controlled work areas (the areas consisting of the work rooms, corridors, etc. that are controlled so as to maintain a uniform quality of cleanliness, and hereinafter referred to as such) where the assembling or packaging operations for the products are conducted, The facilities for supplying water of the quality and quantity necessary for the manufacturing, according to the type of the products and manufacturing process in the work rooms or controlled work areas where the assembling or packaging operations for the products are conducted,23 The sterilisation apparatuses necessary for the manufacturing according to the type of the products, and The facilities and equipment necessary for controlling the sterilisation process according to the type of the products. Validation of Sterilisation Process Article 46 The sterile-medical-device-category manufacturer, etc. shall establish the documented procedure for the validation of the sterilization process. The sterile-medical-device-category manufacturer, etc. shall ensure that the sterilisation24 processes are validated prior to the initial use. The sterile-medical-device-category manufacturer, etc. shall ensure that the records of the validation of each sterilization process are maintained. | No significant differences. The QSR contains a number of prescriptive requirements associated with documentation of validation activities. There are no specific requirements related to process validation of sterilization processes in the QS Reg. The JGMP has specific requirements relating to manufacturing control of sterile medical devices. |
| 7.5.3 Identification and traceability 7.5.3.1 Identification The organization shall identify the product by suitable means throughout product realization, and shall establish documented procedures for such product identification. The organization shall establish documented procedures to ensure that medical devices returned to the organization are identified and distinguished from conforming product [see 6.4 d)]. 7.5.3.2 Traceability 7.5.3.2.1 General The organization shall establish documented procedures for traceability. Such procedures shall define the extent of product traceability and the records required (see 4.2.4, 8.3 and 8.5). Where traceability is a requirement, the organization shall control and record the unique identification of the product (see 4.2.4). NOTE Configuration management is a means by which identification and traceability can be maintained. 7.5.3.2.2 Particular requirements for active implantable medical devices and implantable medical devices In defining the records required for traceability, the organization shall include records of all components, materials and work environment conditions, if these could cause the medical device not to satisfy its specified requirements. The organization shall require that its agents or distributors maintain records of the distribution of medical devices to allow traceability and that such records are available for inspection. Records of the name and address of the shipping package consignee shall be maintained (see 4.2.4). 7.5.3.3 Status identification The organization shall identify the product status with respect to monitoring and measurement requirements. The identification of product status shall be maintained throughout production, storage, installation and servicing of the product to ensure that only product that has passed the required inspections and tests (or released under an authorized concession) is dispatched, used or installed. | 820.60 Identification. Each manufacturer shall establish and maintain procedures for identifying product during all stages of receipt, production, distribution, and installation to prevent mix-ups. 820.65 Traceability. Each manufacturer of a device that is intended for surgical implant into the body or to support or sustain life and whose failure to perform when properly used in accordance with instructions for use provided in the labeling can be reasonably expected to result in a significant injury to the user shall establish and maintain procedures for identifying with a control number each unit, lot, or batch of finished devices and where appropriate components. The procedures shall facilitate corrective action. Such identification shall be documented in the DHR. 820.80 Receiving, in-process, and finished device acceptance (e) Acceptance records. Each manufacturer shall document acceptance activities required by this part. These records shall include: (1) The acceptance activities performed; (2) the dates acceptance activities are performed; (3) the results; (4) the signature of the individual(s) conducting the acceptance activities; and (5) where appropriate the equipment used. These records shall be part of the DHR. 820.86 Acceptance status. Each manufacturer shall identify by suitable means the acceptance status of product, to indicate the conformance or nonconformance of product with acceptance criteria. The identification of acceptance status shall be maintained throughout manufacturing, packaging, labeling, installation, and servicing of the product to ensure that only product which has passed the required acceptance activities is distributed, used, or installed. | Identification Article 47 The manufacturer, etc. shall identify the products by suitable means throughout the product realisation. The manufacturer, etc. shall establish the documented procedure for the product identification specified in preceding Paragraph 1. The manufacturer, etc. shall establish the documented procedure to ensure that the products returned to the manufacturer, etc. are identified and distinguished from the conforming products. (Traceability) Article 48 The manufacturer, etc. shall establish the documented procedure for the traceability. 2. The manufacturer, etc. shall ensure that the documented procedure specified in preceding Paragraph 1 define the extent of the product traceability and the records required. 3. Where the traceability is the requirement, the manufacturer, etc. shall control and record the unique identification of the products. Traceability of Specially Designated Medical Devices Article 49 1. In defining the records required for the traceability, the manufacturer, etc. shall include the records of all the constituent parts, etc. and work environment conditions, if these could cause the products concerned with the specially designated medical devices not to satisfy their specified requirements. The manufacturer, etc. shall ensure that the records of the name and address of the shipping package consignee of the products concerned with the specially designated medical devices are maintained. Status Identification Article 50 The manufacturer, etc. shall identify the product status with respect to the monitoring and measurement requirements. The manufacturer, etc. shall ensure that the identification of the product status is maintained throughout the production, storage, installation and servicing of the products to ensure that only products that have passed the required inspections and tests (or released under an authorised concession (the permission to use or release the products25 that do not conform to the specified requirements after the appropriate verification of the manufacturing control and quality control of the products)) are dispatched, used, or installed. | ISO/DIS13485:2003 and the JGMP addresses the identification and traceability of product returned to the organization in order to ensure it is distinguished from normal product. ISO/DIS 13485:2003 includes some general requirements associated with traceability not found in the QS Reg. The requirements for implantable and active implantable devices found in all three documents are supplementary and all be incorporated into the quality management system of an organization supplying such medical devices. The QSR contains a number of prescriptive requirements associated with the records of acceptance status. |
| 7.5.4 Customer property The organization shall exercise care with customer property while it is under the organization’s control or being used by the organization. The organization shall identify, verify, protect and safeguard customer property provided for use or incorporation into the product. If any customer property is lost, damaged or otherwise found to be unsuitable for use, this shall be reported to the customer and records maintained (see 4.2.4). NOTE Customer property can include intellectual property or confidential health information. | No requirement | Customer Property Article 51 The manufacturer, etc. shall identify, verify, protect and safeguard the customer property provided for the use or incorporation into the products. The manufacturer, etc. shall ensure that, if any property of the customer is lost or otherwise found to be unsuitable for use, the details are reported to the customer and records thereof are maintained. | Aside from the general controls to be exerted by the organization over purchased product, the QSR does not specifically address the issue of care to be exercised over customer property when it is being held or processed by the organization. |
| 7.5.5 Preservation of product The organization shall establish documented procedures or documented work instructions for preserving the conformity of product during internal processing and delivery to the intended destination. This preservation shall include identification, handling, packaging, storage and protection. Preservation shall also apply to the constituent parts of a product. The organization shall establish documented procedures or documented work instructions for the control of product with a limited shelf life or requiring special storage conditions. Such special storage conditions shall be controlled and recorded (see 4.2.4). | 820.120 Device labeling. Each manufacturer shall establish and maintain procedures to control labeling activities. (a) Label integrity. Labels shall be printed and applied so as to remain legible and affixed during the customary conditions of processing, storage, handling, distribution, and where appropriate use. (b) Labeling inspection. Labeling shall not be released for storage or use until a designated individual(s) has examined the labeling for accuracy including, where applicable, the correct expiration date, control number, storage instructions, handling instructions, and any additional processing instructions. The release, including the date and signature of the individual(s) performing the examination, shall be documented in the DHR. (c) Labelling storage. Each manufacturer shall store labeling in a manner that provides proper identification and is designed to prevent mix-ups. (d) Labeling operations. Each manufacturer shall control labeling and packaging operations to prevent labeling mix-ups. The label and labeling used for each production unit, lot, or batch shall be documented in the DHR. e) Control number. Where a control number is required by Sec. 820.65, that control number shall be on or shall accompany the device through distribution. 820.150 Storage. (a) Each manufacturer shall establish and maintain procedures for the control of storage areas and stock rooms for product to prevent mix-ups, damage, deterioration, contamination, or other adverse effects pending use or distribution and to ensure that no obsolete, rejected, or deteriorated product is used or distributed. When the quality of product deteriorates over time, it shall be stored in a manner to facilitate proper stock rotation, and its condition shall be assessed as appropriate. (c) Each manufacturer shall establish and maintain procedures that describe the methods for authorizing receipt from and dispatch to storage areas and stock rooms. 820.130 Device packaging. Each manufacturer shall ensure that device packaging and shipping containers are designed and constructed to protect the device from alteration or damage during the customary conditions of processing, storage, handling, and distribution. 820.140 Handling. Each manufacturer shall establish and maintain procedures to ensure that mix-ups, damage, deterioration, contamination, or other adverse effects to product do not occur during handling. 820.160 Distribution. (a) Each manufacturer shall establish and maintain procedures for control and distribution of finished devices to ensure that only those devices approved for release are distributed and that purchase orders are reviewed to ensure that ambiguities and errors are resolved before devices are released for distribution. Where a device's fitness for use or quality deteriorates over time, the procedures shall ensure that expired devices or devices deteriorated beyond acceptable fitness for use are not distributed. (b) Each manufacturer shall maintain distribution records which include or refer to the location of: (1) The name and address of the initial consignee; (2) The identification and quantity of devices shipped; (3) The date shipped; and (4) Any control number(s) used. | Preservation of Products Article 52 The manufacturer, etc. shall establish the documented procedure or documented work instructions for preserving the conformity of the products (including the identification, handling, packaging, storage and protection) during the internal processing and delivery to the intended destination. The preservation shall also apply to the constituent parts of the products. The manufacturer, etc. shall establish the documented procedure or documented work instructions for the control of the products with the limited shelf life or requiring the special storage conditions. The manufacturer, etc. shall ensure that the special storage conditions specified in preceding Paragraph 2 are controlled and recorded. | While the intent of the three documents is consistent, the QSR contains many prescriptive requirements related to label control and the handling, storage, packaging, preservation, and distribution of product not specifically called out in ISO/DIS 13485:2003 and the JGMP. |
| 7.6 Control of monitoring and measuring devices The organization shall determine the monitoring and measurement to be undertaken and the monitoring and measuring devices needed to provide evidence of conformity of product to determined requirements (see 7.2.1). The organization shall establish documented procedures to ensure that monitoring and measurement can be carried out and are carried out in a manner that is consistent with the monitoring and measurement requirements. Where necessary to ensure valid results, measuring equipment shall a) Be calibrated or verified at specified intervals, or prior to use, against measurement standards traceable to international or national measurement standards; where no such standards exist, the basis used for calibration or verification shall be recorded; b) Be adjusted or re-adjusted as necessary; c) Be identified to enable the calibration status to be determined; d) Be safeguarded from adjustments that would invalidate the measurement result; e) Be protected from damage and deterioration during handling, maintenance and storage. In addition, the organization shall assess and record the validity of the previous measuring results when the equipment is found not to conform to requirements. The organization shall take appropriate action on the equipment and any product affected. Records of the results of calibration and verification shall be maintained (see 4.2.4). When used in the monitoring and measurement of specified requirements, the ability of computer software to satisfy the intended application shall be confirmed. This shall be undertaken prior to initial use and reconfirmed as necessary. NOTE See ISO 10012 for guidance related to measurement management systems. | 820.72 Inspection, measuring, and test equipment. (a) Control of inspection, measuring, and test equipment. Each manufacturer shall ensure that all inspection, measuring, and test equipment, including mechanical, automated, or electronic inspection and test equipment, is suitable for its intended purposes and is capable of producing valid results. Each manufacturer shall establish and maintain procedures to ensure that equipment is routinely calibrated, inspected, checked, and maintained. The procedures shall include provisions for handling, preservation, and storage of equipment, so that its accuracy and fitness for use are maintained. These activities shall be documented. (b) Calibration. Calibration procedures shall include specific directions and limits for accuracy and precision. When accuracy and precision limits are not met, there shall be provisions for remedial action to reestablish the limits and to evaluate whether there was any adverse effect on the device's quality. These activities shall be documented. (1) Calibration standards. Calibration standards used for inspection, measuring, and test equipment shall be traceable to national or international standards. If national or international standards are not practical or available, the manufacturer shall use an independent reproducible standard. If no applicable standard exists, the manufacturer shall establish and maintain an in-house standard. (2) Calibration records. The equipment identification, calibration dates, the individual performing each calibration, and the next calibration date shall be documented. These records shall be displayed on or near each piece of equipment or shall be readily available to the personnel using such equipment and to the individuals responsible for calibrating the equipment. | Control of Monitoring and Measuring Devices Article 53 The manufacturer, etc. shall determine the monitoring and measurement to be undertaken and the monitoring and measuring devices needed to provide the evidence of the conformity of the products to the determined requirements. The manufacturer, etc. shall establish the documented procedure to ensure that the monitoring and measurement specified in preceding Paragraph 1 can be carried out and are carried out in a manner that is consistent with the monitoring and measurement requirements. Where necessary to ensure valid results, the manufacturer, etc. shall ensure that the measuring equipment meets the following requirements. To be calibrated or verified at the specified intervals, or prior to the use, against the measurement standards traceable to the measurement standards (in case where no such standards exist, the basis used for the calibration or verification shall be recorded), To be adjusted or re-adjusted as necessary,26 To be identified to enable the calibration status to be determined, To be safeguarded from the adjustments that would invalidate the measurement result, and To be protected from damage and deterioration during the handling, maintenance and storage. The manufacturer, etc. shall assess and record the validity of the previous measuring results when the equipment is found not to conform to the requirements. The manufacturer, etc., in case where they correspond to the case specified in preceding Paragraph 4, shall take appropriate action on the equipment and for any of the products affected. The manufacturer, etc. shall ensure that the records of the results of the calibration and verification are maintained. The manufacturer, etc. shall ensure that, when used in the monitoring and measurement of the specified requirements, the ability of the computer software to satisfy the intended application is confirmed prior to the initial use and reconfirmed as necessary. | The intent of the three documents is consistent, with ISO/DIS 13485:2003 and the JGMP giving generalized guidance as to the control of monitoring and measuring devices and the QSR focusing more specifically on the process of calibration of such equipment. |
| 8 Measurement, analysis and improvement 8.1 General The organization shall plan and implement the monitoring, measurement, analysis and improvement processes needed a) To demonstrate conformity of the product, b) To ensure conformity of the quality management system, and c) To maintain the effectiveness of the quality management system. This shall include determination of applicable methods, including statistical techniques, and the extent of their use. NOTE National or regional regulations might require documented procedures for implementation and control of the application of statistical techniques. | 820.250 Statistical techniques. (a) Where appropriate, each manufacturer shall establish and maintain procedures for identifying valid statistical techniques required for establishing, controlling, and verifying the acceptability of process capability and product characteristics. (b) Sampling plans, when used, shall be written and based on a valid statistical rationale. Each manufacturer shall establish and maintain procedures to ensure that sampling methods are adequate for their intended use and to ensure that when changes occur the sampling plans are reviewed. These activities shall be documented. | Section 6 Measurement, Analysis and Improvement Measurement, Analysis and Improvement Article 54 The manufacturer, etc. shall establish a plan (including determination of the applicable testing procedure, such as the statistical techniques, and the extent of the application) and implement the monitoring, measurement, analysis and improvement processes needed for the following duties. To demonstrate the conformity of the products, and To ensure the conformity of the quality management system, and to maintain the effectiveness of the quality management system. | There is no one section of the QSR that corresponds to this clause of ISO/DIS 13485:2003 and the JGMP. It is clear from a reading of the overall QSR that the objectives of this clause of ISO/DIS 13485:2003 and the JGMP, and the QSR are consistent. |
| 8.2 Monitoring and measurement 8.2.1 Feedback As one of the measurements of the performance of the quality management system, the organization shall monitor information relating to whether the organization has met customer requirements. The methods for obtaining and using this information shall be determined. The organization shall establish a documented procedure for a feedback system [see 7.2.3 c)] to provide early warning of quality problems and for input into the corrective and preventive action processes (see 8.5.2 and 8.5.3). If national or regional regulations require the organization to gain experience from the post-production phase, the review of this experience shall form part of the feedback system (see 8.5.1). | 820.198 Complaint files. (a) Each manufacturer shall maintain complaint files. Each manufacturer shall establish and maintain procedures for receiving, reviewing, and evaluating complaints by a formally designated unit. Such procedures shall ensure that: (1) All complaints are processed in a uniform and timely manner; (2) Oral complaints are documented upon receipt; and (3) Complaints are evaluated to determine whether the complaint represents an event, which is required to be reported to FDA under part 803 or 804 of this chapter, Medical Device Reporting. (b) Each manufacturer shall review and evaluate all complaints to determine whether an investigation is necessary. When no investigation is made, the manufacturer shall maintain a record that includes the reason no investigation was made and the name of the individual responsible for the decision not to investigate. (c) Any complaint involving the possible failure of a device, labeling, or packaging to meet any of its specifications shall be reviewed, evaluated, and investigated, unless such investigation has already been performed for a similar complaint and another investigation is not necessary. (d) Any complaint that represents an event which must be reported to FDA under part 803 or 804 of this chapter shall be promptly reviewed, evaluated, and investigated by a designated individual(s) and shall be maintained in a separate portion of the complaint files or otherwise clearly identified. In addition to the information required by Sec. 820.198(e), records of investigation under this paragraph shall include a determination of: (1) Whether the device failed to meet specifications; (2) Whether the device was being used for treatment or diagnosis; and (3) The relationship, if any, of the device to the reported incident or adverse event. (e) When an investigation is made under this section, a record of the investigation shall be maintained by the formally designated unit identified in paragraph (a) of this section. The record of investigation shall include: (1) The name of the device; (2) The date the complaint was received; (3) Any device identification(s) and control number(s) used; (4) The name, address, and phone number of the complainant; (5) The nature and details of the complaint; (6) The dates and results of the investigation; (7) Any corrective action taken; and (8) Any reply to the complainant. (f) When the manufacturer's formally designated complaint unit is located at a site separate from the manufacturing establishment, the investigated complaint(s) and the record(s) of investigation shall be reasonably accessible to the manufacturing establishment. (g) If a manufacturer's formally designated complaint unit is located outside of the United States, records required by this section shall be reasonably accessible in the United States at either: (1) A location in the United States where the manufacturer's records are regularly kept; or (2) The location of the initial distributor. | Feedback Article 55 As one of the measurements of the performance of the quality management system, the manufacturer, etc. shall monitor the information relating to whether the manufacturer, etc. have met the customer requirements. The manufacturer, etc. shall ensure that the methods for obtaining and using the information specified in preceding Paragraph 1 are determined. The manufacturer, etc. shall establish the documented procedure for the feedback system to provide early warning of the quality problems and for the input into the corrective and preventive action processes.27 The manufacturer, etc. shall ensure that the review of the experience gained from the post-production phase forms part of the feedback system specified in preceding Paragraph 3. | The QSR contains prescriptive requirements regarding the handling of complaints and other user input that would be useful in conducting corrective or preventive action. These requirements are not spelled out in detail in ISO/DIS 13485:2003 and the JGMP, but compliance with the requirements set out in the QSR would satisfy the requirements of ISO/DIS 13485:2003 and the JGMP. |
| 8.2.2 Internal audit The organization shall conduct internal audits at planned intervals to determine whether the quality management system a) Conforms to the planned arrangements (see 7.1), to the requirements of this International Standard and to the quality management system requirements established by the organization, and b) Is effectively implemented and maintained. An audit programme shall be planned, taking into consideration the status and importance of the processes and areas to be audited, as well as the results of previous audits. The audit criteria, scope, frequency and methods shall be defined. Selection of auditors and conduct of audits shall ensure objectivity and impartiality of the audit process. Auditors shall not audit their own work. The responsibilities and requirements for planning and conducting audits, and for reporting results and maintaining records (see 4.2.4) shall be defined in a documented procedure. The management responsible for the area being audited shall ensure that actions are taken without undue delay to eliminate detected nonconformities and their causes. Follow-up activities shall include the verification of the actions taken and the reporting of verification results (see 8.5.2). NOTE See ISO 19011 for guidance related to quality auditing. | 820.22 Quality audit. Each manufacturer shall establish procedures for quality audits and conduct such audits to assure that the quality system is in compliance with the established quality system requirements and to determine the effectiveness of the quality system. Quality audits shall be conducted by individuals who do not have direct responsibility for the matters being audited. Corrective action(s), including a reaudit of deficient matters, shall be taken when necessary. A report of the results of each quality audit, and reaudit(s) where taken, shall be made and such reports shall be reviewed by management having responsibility for the matters audited. The dates and results of quality audits and reaudits shall be documented. | Internal Audit Article 56 The manufacturer, etc. shall conduct the internal audits at planned intervals to determine whether the quality management system meets the following requirements. To conform to the arrangements defined in the product realisation plan, to the requirements of this Ministerial Ordinance and to the quality management system requirements established by the manufacturer, etc., and To be effectively implemented and maintained. The manufacturer, etc. shall ensure that the audit programme is planned, taking into consideration the status and importance of the processes and areas to be audited, as well as the results of the previous audits. The manufacturer, etc. shall ensure that the audit criteria, scope, frequency and methods are defined. The manufacturer, etc. shall ensure that the selection of the personnel who conduct the audits (hereinafter referred to as “auditors”) and the conduct of the audits ensure objectivity and impartiality of the audit process. The manufacturer, etc. shall ensure that the auditors do not audit their own work. The manufacturer, etc. shall ensure that the responsibilities and requirements for planning and conducting the audits, and for reporting the results and maintaining the records are defined in the documented procedure. The manufacturer, etc. shall ensure that the management responsible for the area being audited ensures that the actions are taken without undue delay to eliminate the detected nonconformities and their causes and the follow-up activities include the verification of the actions taken and the reporting of the verification results. | No significant differences. |
| 8.2.3 Monitoring and measurement of processes The organization shall apply suitable methods for monitoring and, where applicable, measurement of the quality management system processes. These methods shall demonstrate the ability of the processes to achieve planned results. When planned results are not achieved, correction and corrective action shall be taken, as appropriate, to ensure conformity of the product. | 820.70 Production and process controls. (a) General. Each manufacturer shall develop, conduct, control, and monitor production processes to ensure that a device conforms to its specifications. Where deviations from device specifications could occur as a result of the manufacturing process, the manufacturer shall establish and maintain process control procedures that describe any process controls necessary to ensure conformance to specifications. Where process controls are needed they shall include: 1) Documented instructions, standard operating procedures (SOP's), and methods that define and control the manner of production; (2) Monitoring and control of process parameters and component and device characteristics during production; (3) Compliance with specified reference standards or codes; (4) The approval of processes and process equipment; and (5) Criteria for workmanship, which shall be expressed in documented standards or by means of, identified and approved representative samples. 820.250 Statistical techniques. (a) Where appropriate, each manufacturer shall establish and maintain procedures for identifying valid statistical techniques required for establishing, controlling, and verifying the acceptability of process capability and product characteristics. (b) Sampling plans, when used, shall be written and based on a valid statistical rationale. Each manufacturer shall establish and maintain procedures to ensure that sampling methods are adequate for their intended use and to ensure that when changes occur the sampling plans are reviewed. These activities shall be documented. | Monitoring and Measurement of Processes Article 57 The manufacturer, etc. shall apply suitable methods for monitoring and, where applicable, measurement of the quality management system processes. The manufacturer, etc. shall ensure that the methods specified in preceding Paragraph 1 demonstrate the ability of the processes to achieve the results defined in the plan specified in Paragraph 1 of Article 14. The manufacturer, etc. shall ensure that, when the planned results specified in Paragraph 1 of Article 14 are not achieved, the correction and corrective action shall be taken, as appropriate, to ensure the conformity of the products.28 Monitoring and Measurement of Specially Designated Medical Devices Article 59 The manufacturer, etc. shall record the identity of the personnel performing any inspection or testing specified in preceding Article of the products concerned with the specially designated medical devices. | The intent of the three documents is consistent, even though the QSR is far more detailed and has a product focus. |
| 8.2.4 Monitoring and measurement of product 8.2.4.1 General requirements The organization shall monitor and measure the characteristics of the product to verify that product requirements have been met. This shall be carried out at appropriate stages of the product realization process in accordance with the planned arrangements (see 7.1) and documented procedures (see 7.5.1.1). Evidence of conformity with the acceptance criteria shall be maintained. Records shall indicate the person(s) authorizing release of product (see 4.2.4). Product release and service delivery shall not proceed until the planned arrangements (see 7.1) have been satisfactorily completed. 8.2.4.2 Particular requirement for active implantable medical devices and implantable medical devices The organization shall record (see 4.2.4) the identity of personnel performing any inspection or testing | Receiving, in-process, and finished device acceptance (a) General. Each manufacturer shall establish and maintain procedures for acceptance activities. Acceptance activities include inspections, tests, or other verification activities. (b) Receiving acceptance activities. Each manufacturer shall establish and maintain procedures for acceptance of incoming product. Incoming product shall be inspected, tested, or otherwise verified as conforming to specified requirements. Acceptance or rejection shall be documented. (c) In-process acceptance activities. Each manufacturer shall establish and maintain acceptance procedures, where appropriate, to ensure that specified requirements for in-process product are met. Such procedures shall ensure that in-process product is controlled until the required inspection and tests or other verification activities have been completed, or necessary approvals are received, and are documented. (d) Final acceptance activities. Each manufacturer shall establish and maintain procedures for finished device acceptance to ensure that each production run, lot, or batch of finished devices meets acceptance criteria. Finished devices shall be held in quarantine or otherwise adequately controlled until released. Finished devices shall not be released for distribution until: (1) The activities required in the DMR are completed; (2) The associated data and documentation is reviewed; (3) The release is authorized by the signature of a designated individual(s); and (4) The authorization is dated. (e) Acceptance records. Each manufacturer shall document acceptance activities required by this part. These records shall include: (1) The acceptance activities performed; (2) the dates acceptance activities are performed; (3) the results; (4) the signature of the individual(s) conducting the acceptance activities; and (5) where appropriate the equipment used. These records shall be part of the DHR. 820.250 Statistical techniques. (a) Where appropriate, each manufacturer shall establish and maintain procedures for identifying valid statistical techniques required for establishing, controlling, and verifying the acceptability of process capability and product characteristics. (b) Sampling plans, when used, shall be written and based on a valid statistical rationale. Each manufacturer shall establish and maintain procedures to ensure that sampling methods are adequate for their intended use and to ensure that when changes occur the sampling plans are reviewed. These activities shall be documented. | Monitoring and Measurement of Products Article 58 The manufacturer, etc. shall monitor and measure the characteristics of the products to verify that the product requirements have been met. The manufacturer, etc. shall ensure that the monitoring and measurement specified in preceding Paragraph 1 should be carried out at appropriate stages of the product realisation process in accordance with the planned arrangements and documented procedure specified in Item (2) of Paragraph 1 of Article 41. The manufacturer, etc. shall ensure that the evidence of the conformity with the acceptance criteria is maintained. The manufacturer, etc. shall ensure that the records indicate the personnel authorising the release of the products. The manufacturer, etc. shall ensure that the product release and service delivery do not proceed until the arrangements defined in the product realisation plan have been satisfactorily completed. | The intent of the three documents is consistent, even though the QSR is far more detailed and prescriptive. The acceptance records requirements of the QSR would appear to satisfy the implantable and active implantable requirements of ISO/DIS 13485:2003 and the JGMP. |
| 8.3 Control of nonconforming product The organization shall ensure that product which does not conform to product requirements is identified and controlled to prevent its unintended use or delivery. The controls and related responsibilities and authorities for dealing with nonconforming product shall be defined in a documented procedure. The organization shall deal with nonconforming product by one or more of the following ways: a) By taking action to eliminate the detected nonconformity; b) By authorizing its use, release or acceptance under concession; c) By taking action to preclude its original intended use or application. The organization shall ensure that nonconforming product is accepted by concession only if regulatory requirements are met. Records of the identity of the person(s) authorizing the concession shall be maintained (see 4.2.4). Records of the nature of nonconformities and any subsequent actions taken, including concessions obtained, shall be maintained (see 4.2.4). When nonconforming product is corrected it shall be subject to re-verification to demonstrate conformity to the requirements. When nonconforming product is detected after delivery or use has started, the organization shall take action appropriate to the effects, or potential effects, of the nonconformity. If product needs to be reworked (one or more times), the organization shall document the rework process in a work instruction that has undergone the same authorization and approval procedure as the original work instruction. Prior to authorization and approval of the work instruction, a determination of any adverse effect of the rework upon product shall be made and documented (see 4.2.3 and 7.5.1). | 820.90 Nonconforming product. (a) Control of nonconforming product. Each manufacturer shall establish and maintain procedures to control product that does not conform to specified requirements. The procedures shall address the identification, documentation, evaluation, segregation, and disposition of nonconforming product. The evaluation of nonconformance shall include a determination of the need for an investigation and notification of the persons or organizations responsible for the nonconformance. The evaluation and any investigation shall be documented. (b) Nonconformity review and disposition. (1) Each manufacturer shall establish and maintain procedures that define the responsibility for review and the authority for the disposition of nonconforming product. The procedures shall set forth the review and disposition process. Disposition of nonconforming product shall be documented. Documentation shall include the justification for use of nonconforming product and the signature of the individual(s) authorizing the use. (2) Each manufacturer shall establish and maintain procedures for rework; to include retesting and reevaluation of the nonconforming product after rework, to ensure that the product meets its current approved specifications. Rework and reevaluation activities, including a determination of any adverse effect from the rework upon the product, shall be documented in the DHR. | Control of Nonconforming Products Article 60 The manufacturer, etc. shall ensure that the products which do not conform to the product requirements (hereinafter referred to as “nonconforming products”) are identified and controlled to prevent their unintended use or delivery. The manufacturer, etc. shall ensure that the controls and related responsibilities and authorities for dealing with the nonconforming products are defined in the documented procedure. The manufacturer, etc. shall deal with the nonconforming products by one or more of the following ways. To take actions to eliminate the detected nonconformity, To authorise their use, release or acceptance under the concession, or To take actions to preclude their original intended use or application. The manufacturer, etc. shall ensure that the nonconforming products are accepted by the concession only if the regulatory requirements are met.29 The manufacturer, etc. shall ensure that the records of the identity of the personnel authorising the concession are maintained. The manufacturer, etc. shall ensure that the records of the nature of the nonconformities and any subsequent actions taken, including the concessions obtained, are maintained. The manufacturer, etc. shall ensure, when the nonconforming products are corrected, that they are subject to re-verification to demonstrate the conformity to the requirements. When the nonconforming products are detected after the delivery or use has started, the manufacturer, etc. shall take actions appropriate to the effects, or potential effects, of the nonconformity. If the products need to be reworked (one or more times), the manufacturer, etc. shall document the rework process in the work instructions that has undergone the same authorisation and approval procedure as the original work instructions. The manufacturer, etc. shall ensure that, prior to the authorisation and approval of the work instructions, a determination of any adverse effect of the rework upon the products is made and documented. | The intent of the three documents is consistent. The QSR provides more detail as to the items to be recorded in a nonconforming product situation. It explicitly addresses the subject of the need for an investigation in such a situation. ISO/DIS 13485:2003 and the JGMP address the handling of already released product that is subject to a finding of nonconformity. ISO/DIS 13485:2003 and the JGMP provide requirements related to rework of non-conforming product. |
| 8.4 Analysis of data The organization shall establish documented procedures to determine, collect and analyse appropriate data to demonstrate the suitability and effectiveness of the quality management system and to evaluate if improvement of the effectiveness of the quality management system can be made. This shall include data generated as a result of monitoring and measurement and from other relevant sources. The analysis of data shall provide information relating to a) Feedback (see 8.2.1), b) Conformity to product requirements (see 7.2.1), c) Characteristics and trends of processes and products including opportunities for preventive action, and d) Suppliers. Records of the results of the analysis of data shall be maintained (see 4.2.4). | 820.250 Statistical techniques. (a) Where appropriate, each manufacturer shall establish and maintain procedures for identifying valid statistical techniques required for establishing, controlling, and verifying the acceptability of process capability and product characteristics. (b) Sampling plans, when used, shall be written and based on a valid statistical rationale. Each manufacturer shall establish and maintain procedures to ensure that sampling methods are adequate for their intended use and to ensure that when changes occur the sampling plans are reviewed. These activities shall be documented. | Analysis of Data Article 61 The manufacturer, etc. shall establish the documented procedure to determine, collect and analyse appropriate data (including the data generated as the result of the monitoring and measurement and from other relevant sources) to demonstrate the suitability and effectiveness of the quality management system and to evaluate if improvement of the effectiveness of the quality management system can be made. The manufacturer, etc. shall ensure that the analysis of the data specified in preceding Paragraph 1 provides the information relating to the following items. The feedback collected in accordance with the documented procedure specified in Paragraph 3 of Article 55, The conformity to the product requirements, The characteristics and trends of the processes and products including the opportunities for the preventive actions, and The suppliers. The manufacturer, etc. shall ensure that the records of the results of the analysis of the data specified in preceding two Paragraphs 1 and 2 are maintained. | There is no one section of the QSR that corresponds directly with sub clause 8.4 of ISO/DIS 13485:2003 and article 61 of the JGMP. Requirements for analysis of data related to various elements of the quality management systems are contained in a number of QSR sections. The intent of the three documents as it relates to analysis of data is consistent. |
| 8.5 Improvement 8.5.1 General The organization shall identify and implement any changes necessary to ensure and maintain the continued suitability and effectiveness of the quality management system through the use of the quality policy, quality objectives, audit results, analysis of data, corrective and preventive actions and management review. The organization shall establish documented procedures for the issue and implementation of advisory notices. These procedures shall be capable of being implemented at any time. Records of all customer complaint investigations shall be maintained (see 4.2.4). If investigation determines that the activities outside the organization contributed to the customer complaint, relevant information shall be exchanged between the organizations involved (see 4.1). If any customer complaint is not followed by corrective and/or preventive action, the reason shall be authorized (see 5.5.1) and recorded (see 4.2.4). If national or regional regulations require notification of adverse events that meet specified reporting criteria, the organization shall establish documented procedures to such notification to regulatory authorities. | Management responsibilities (c) Management review. Management with executive responsibility shall review the suitability and effectiveness of the quality system at defined intervals and with sufficient frequency according to established procedures to ensure that the quality system satisfies the requirements of this part and the manufacturer's established quality policy and objectives. The dates and results of quality system reviews shall be documented. 820.198 Complaint files. (a) Each manufacturer shall maintain complaint files. Each manufacturer shall establish and maintain procedures for receiving, reviewing, and evaluating complaints by a formally designated unit. Such procedures shall ensure that: (1) All complaints are processed in a uniform and timely manner; (2) Oral complaints are documented upon receipt; and (3) Complaints are evaluated to determine whether the complaint represents an event, which is required to be reported to FDA under part 803 or 804 of this chapter, Medical Device Reporting. (b) Each manufacturer shall review and evaluate all complaints to determine whether an investigation is necessary. When no investigation is made, the manufacturer shall maintain a record that includes the reason no investigation was made and the name of the individual responsible for the decision not to investigate. (c) Any complaint involving the possible failure of a device, labeling, or packaging to meet any of its specifications shall be reviewed, evaluated, and investigated, unless such investigation has already been performed for a similar complaint and another investigation is not necessary. (d) Any complaint that represents an event which must be reported to FDA under part 803 or 804 of this chapter shall be promptly reviewed, evaluated, and investigated by a designated individual(s) and shall be maintained in a separate portion of the complaint files or otherwise clearly identified. In addition to the information required by Sec. 820.198(e), records of investigation under this paragraph shall include a determination of: (1) Whether the device failed to meet specifications; (2) Whether the device was being used for treatment or diagnosis; and (3) The relationship, if any, of the device to the reported incident or adverse event. (e) When an investigation is made under this section, a record of the investigation shall be maintained by the formally designated unit identified in paragraph (a) of this section. The record of investigation shall include: (1) The name of the device; (2) The date the complaint was received; (3) Any device identification(s) and control number(s) used; (4) The name, address, and phone number of the complainant; (5) The nature and details of the complaint; (6) The dates and results of the investigation; (7) Any corrective action taken; and (8) Any reply to the complainant. (f) When the manufacturer's formally designated complaint unit is located at a site separate from the manufacturing establishment, the investigated complaint(s) and the record(s) of investigation shall be reasonably accessible to the manufacturing establishment. (g) If a manufacturer's formally designated complaint unit is located outside of the United States, records required by this section shall be reasonably accessible in the United States at either: (1) A location in the United States where the manufacturer's records are regularly kept; or (2) The location of the initial distributor. | Improvement Article 62 The manufacturer, etc. shall identify and implement any changes necessary30 to ensure and maintain the continued suitability and effectiveness of the quality management system through the use of the quality policy, quality objectives, audit results, analysis of the data, corrective and preventive actions and management review. The manufacturer, etc. shall establish the documented procedure for the issue and implementation of the advisory notices, and shall ensure that such a procedure is capable of being implemented at any time, with the proviso that this provision shall not apply in case where the customer issues and implements the advisory notices, and the manufacturer, etc. shall provide the customer with the information necessary for issuing the advisory notices). The manufacturer, etc. shall ensure that the records of all the customer compliant investigations are maintained. If the investigation specified in preceding Paragraph 3 determines that the activities outside the manufacturer, etc. contributed to the customer compliant, the manufacturer, etc. shall ensure that the relevant information is exchanged between the organisations involved. The manufacturer, etc. shall ensure, if any customer compliant is not followed by the corrective and/or preventive action, that the reason is authorised and recorded. The manufacturer, etc. shall, in case where they become aware of the information on the products specified in Paragraph 2 of Article 253 of Enforcement Regulations, establish the documented procedure to notify such information to the customers. | There is no one section of the QSR that corresponds directly with sub clause 8.5.1 of ISO/DIS 13485:2003 and the JGMP article 62. The intent of the three documents as it relates to improvement of the quality management system through the use of corrective and preventive action is consistent. The requirements of the QSR are significantly more prescriptive related to the handling and documentation of complaints. |
| 8.5.2 Corrective action The organization shall take action to eliminate the cause of nonconformities in order to prevent recurrence. Corrective actions shall be appropriate to the effects of the nonconformities encountered. A documented procedure shall be established to define requirements for a) Reviewing nonconformities (including customer complaints), b) Determining the causes of nonconformities, c) Evaluating the need for action to ensure that nonconformities do not recur, d) Determining and implementing action needed, including, if appropriate, updating documentation (see 4.2), e) Recording of the results of any investigation and of action taken (see 4.2.4), and f) Reviewing the corrective action taken and its effectiveness. | 820.100 Corrective and preventive action. (a) Each manufacturer shall establish and maintain procedures for implementing corrective and preventive action. The procedures shall include requirements for: (1) Analyzing processes, work operations, concessions, quality audit reports, quality records, service records, complaints, returned product, and other sources of quality data to identify existing and potential causes of nonconforming product, or other quality problems. Appropriate statistical methodology shall be employed where necessary to detect recurring quality problems; (2) Investigating the cause of nonconformities relating to product, processes, and the quality system; (3) Identifying the action(s) needed to correct and prevent recurrence of nonconforming product and other quality problems; (4) Verifying or validating the corrective and preventive action to ensure that such action is effective and does not adversely affect the finished device; (5) Implementing and recording changes in methods and procedures needed to correct and prevent identified quality problems; (6) Ensuring that information related to quality problems or nonconforming product is disseminated to those directly responsible for assuring the quality of such product or the prevention of such problems; and (7) Submitting relevant information on identified quality problems, as well as corrective and preventive actions, for management review. (b) All activities required under this section, and their results, shall be documented. | Corrective Actions Article 63 The manufacturer, etc. shall take the corrective actions which are appropriate to the effects of the nonconformities encountered to eliminate the cause of the nonconformities in order to prevent the recurrence. The manufacturer, etc. shall ensure that the documented procedure is established to define the following requirements. To review the nonconformities (including the customer complaints), To determine the causes of the nonconformities, To evaluate the need for the actions to ensure that the nonconformities do not recur, To determine and implement the actions needed, including, if appropriate, updating the documentation, To record the results of any investigation and of the actions taken, and To review the corrective actions taken and their effectiveness.31 | The intent of the three documents is consistent. The requirements as set out by the QSR are more prescriptive. |
| 8.5.3 Preventive action The organization shall determine action to eliminate the causes of potential nonconformities in order to prevent their occurrence. Preventive actions shall be appropriate to the effects of the potential problems. A documented procedure shall be established to define requirements for a) Determining potential nonconformities and their causes, b) Evaluating the need for action to prevent occurrence of nonconformities, c) Determining and implementing action needed, d) Recording of the results of any investigations and of action taken (see 4.2.4), and e) Reviewing preventive action taken and its effectiveness | 820.100 Corrective and preventive action. (a) Each manufacturer shall establish and maintain procedures for implementing corrective and preventive action. The procedures shall include requirements for: (1) Analyzing processes, work operations, concessions, quality audit reports, quality records, service records, complaints, returned product, and other sources of quality data to identify existing and potential causes of nonconforming product, or other quality problems. Appropriate statistical methodology shall be employed where necessary to detect recurring quality problems; (2) Investigating the cause of nonconformities relating to product, processes, and the quality system; (3) Identifying the action(s) needed to correct and prevent recurrence of nonconforming product and other quality problems; (4) Verifying or validating the corrective and preventive action to ensure that such action is effective and does not adversely affect the finished device; (5) Implementing and recording changes in methods and procedures needed to correct and prevent identified quality problems; (6) Ensuring that information related to quality problems or nonconforming product is disseminated to those directly responsible for assuring the quality of such product or the prevention of such problems; and (7) Submitting relevant information on identified quality problems, as well as corrective and preventive actions, for management review. (b) All activities required under this section, and their results, shall be documented. | Preventive Actions Article 64 The manufacturer, etc. shall determine the preventive actions which are appropriate to the effects of the potential problems to eliminate the causes of the potential nonconformities in order to prevent their occurrence. The manufacturer, etc. shall ensure that the documented procedure is established to define the following requirements. To determine the potential nonconformities and their causes, To evaluate the need for the actions to prevent the occurrence of the nonconformities, To determine and implement the actions needed, To record the results of any investigations and of the actions taken, and To review the preventive actions taken and their effectiveness. | The intent of the three documents is consistent. The requirements as set out by the QSR are more prescriptive. |