(COMPANY) LABORATORY PROFICIENCY REPORT

11/6/00

| ISO / IEC Guide 25 Requirements | Current Activities (What are we doing to address this) | Analysis of Results (What is this telling us?) | Conclusions (Are we OK?) | Future Plans (including responsibilities and timelines) |
| --- | --- | --- | --- | --- |
| a) Statistical Analysis of data | None | N/A | N/A | Will assess the value of charting the weekly micrometer check results in X-R charts. Torque tests may be more appropriate for statistical analysis (charting). Responsibility: ? Due date: 3/1/01 |
| b) Inter-laboratory comparisons | Sent Outside Micrometer to Quality Calibration Service, Inc. (A2LA accredited ) for comparison of results | No significant variation | Results provide confidence in our calibrations using gage blocks | Calipers will be sent out to another lab as next check of lab proficiency using gage blocks. Responsibility: ? Due date: 3/1/01 Torque wrench will be purchased and sent out to another lab as next check of lab proficiency for torque. Responsibility: ? Due date: 3/1/01 Will determine what other calibration categories in the lab scope should be included in inter-laboratory comparisons. |
| c) regular use of certified reference materials and/or checks using secondary reference materials | This method of proficiency testing is not used at this time. | N/A | N/A | Will continue to assess the value of this method. |
| d) replicate testing | Nothing to date | N/A | N/A | Gage R&R studies will be conducted on torque wrenches. Responsibility: ? Due date: 3/1/01 |
| e) re-testing of retained units | Outside micrometers (C-0053) re-tested on 9/18/00 and 9/26/00. | No significant variation. | Results provide confidence in our calibrations using gage blocks | Will continue to re-test micrometer (C-0053) on a weekly basis. Will assess the value of charting the results in X-R charts. Responsibility: ? Due date: weekly |
| f) correlation of results using different characteristics of a unit | This method of proficiency testing is not used at this time. | N/A | N/A | Will continue to assess the value of this method. |