14	STORAGE, DISTRIBUTION, AND INSTALLATION





�
INTRODUCTION





	The device Quality System (QS) regulation covers the manufacture, storage (820.150), distribution (820.160) and installation (820.170) of finished devices. For manufacturers and importers, distribution is one of the most important steps in their quality system. After a product is distributed, a manufacturer rarely has direct control over the product or how it is used. Thus, it is important that controls be in place to assure that only correctly labeled, packaged and approved finished devices are distributed and, if necessary, installed. 





Holding and Distribution Procedures





	Section 820.160 requires that the purchase order be reviewed to ensure that ambiguities and errors are resolved before devices are released for distribution. Manufacturers should have a program to reduce problems in this area. Marketing personnel should be adequately trained. Sales specification flyers and catalogs should be carefully written and kept current to reduce ordering problems. Incoming purchase orders should be checked and ambiguities and errors resolved. After receipt by appropriately trained personnel, orders should be reviewed immediately. If purchase orders are reviewed late in the manufacturing process or just before distribution, the value of the review may be significantly reduced. Where the customer includes specifications, each specified parameter should be checked against the corresponding parameter for the device. A checklist of device parameters may be a helpful tool for this review and should be filed with, or keyed to, the purchase order. 





	The QS regulation (820.60), Identification, requires manufacturers to set up and maintain identity control of their products from component receipt, production, distribution and through installation to prevent mixups. The regulation also requires that written procedures be provided for control and distribution of finished devices (820.160). The purpose of this requirement is to assure that only approved devices are distributed. Each manufacturer should determine what written procedures are needed to assure that only ìapproved for releaseî devices are distributed from the manufacturer. If a manufacturer believes written procedures will not contribute to assuring that only ìapproved for releaseî devices are distributed by their manufacturer, they should be able to defend their decision. For example, the control is integrated into the activities required to package the device or to complete the device history record. This flexibility is allowed by section 820.5, Quality system, of the QS regulation.





	Many manufacturers mark their released finished devices or identify them by location or packaging so that a simple visual check is sufficient to indicate whether the product is acceptable to release for distribution. For example, radiation-emitting electronic products are subject to a performance standard. The application of the certification label is often the last step in approving product release for distribution, and this label is used to distinguish such devices. After final release, the crating of large equipment is a very distinguishing feature. These types of operations may preclude the need for a separate written procedure.





	For interstate contract sterilization, 21 C.F.R. section 801.150(e) requires a written agreement between the parties which details the necessary procedures to help prevent the erroneous release of packaged and labeled ìsterileî but not yet sterilized devices that appear to be, but are not, ready for release. Regardless of whether 801.150(e) applies, the QS regulation requires controls, as necessary, to prevent mixups in complex situations such as contract sterilization. For consistency, a contract as described by 801.150(e) is commonly used by manufacturers for interstate and intrastate shipments. Compliance with such a contract satisfies the applicable GMP requirements. (See Chapter 10, Purchasing and Acceptance Activities, and Compliance Policy Guide 7382.830B for details.)





	Sometimes manufacturers need to ship ìfinished devicesî that have not been officially released because the final test data is not yet available. The critical factor is that the device still remains under the manufacturerís control. The most common example occurs when a manufacturer is waiting for the results from biological indicator tests. FDA permits manufacturers to ship such devices under quarantine to their own controlled warehouses where the devices may be readily recalled prior to any use, if the need arises. Manufacturers should not ship non-released devices to routine distributors or anyone outside of their direct control. Non-released products or products on ìholdî for any quality reason should be controlled to prevent release. A suitable control is quarantine with a label on the units, pallets, etc., to indicate their status.





Warehouse Storage





	Storage should always be done under systematic, orderly conditions (820.150). Manufacturers should use a first-in, first-out (FIFO) distribution system when fitness for use of a device deteriorates over time (820.150). 





	When a controlled environment is necessary to prevent abnormal deterioration, the environment should be specified, controlled, and monitored according to sections 820.70(c) (see Chapter 6, Buildings and Environment). Environmental specifications, such as storage temperature, should be included in the device master record. 





	The storage and handling of devices to be distributed may involve extensive activities (820.140 and 820.150). For example, damaged, recalled or returned devices should be suitably marked and segregated from devices acceptable for release (820.86). Returned devices should be handled and stored such that the cause of failure or other useful information is not destroyed. Returned defective devices should be formally investigated according to 820.100 Corrective and Preventive Action and any associated complaints investigated according to 820.198. Therefore, manufacturers will need controls to assure that returned defective devices do not dead-end in the warehouse, but are expeditiously routed to the appropriate department for evaluation, investigation, conclusions and follow-up (see Chapter 15, Complaint Files).





Distribution Records





	Quality System section 820.184, Device History Record (DHR), requires manufacturers of devices to maintain basic records for:





∫	dates of manufacture, 


∫	the quantity manufactured, 


∫	quantity released for distribution,


∫	acceptance records, 


∫	primary identification labels and labeling used, and


∫	any device identification and control number used. 





Section 820.160, Distribution, requires the following records:





∫	name and address of the initial consignee,


∫	identification and quantity of devices shipped,


∫	date shipped, and 


∫	any control number used. 





	Some of the above information necessary for the distribution records is a duplicate of Device History Record (DHR) requirements. These duplications may be copied or transferred electronically from the DHR. If appropriate, a manufacturer may combine the records by adding the distribution information to the DHR.





	In addition to the above requirements, manufacturers of implantable devices and life sustaining devices, the failure of which during use could result in significant injury to the user, are required to establish and maintain procedures for identifying with a control number each unit, lot, or batch of finished devices and, where appropriate, components (820.65). A partial list of traceable devices that meet this definition is printed at the end of this chapter.





	Distribution records may be the same as, or part of, the normal business records. Generation of a separate record is not required unless the business records are not readily available, e.g., not maintained at the same establishment as the device history record and not readily retrievable electronically. Many manufacturers, also keep distribution records for billing and market survey purposes. 





	Manufacturers of radiological electronic products listed in 21 CFR 1002.1, Record and Reporting Requirements By Product, shall maintain distribution records that will enable them to trace specific products or production lots to distributors, or to dealers in those instances in which the manufacturer distributes directly to dealers (See 21 CFR 1002.30, Records to be Maintained by Manufacturers).





	Distribution records shall be kept for a period of time equivalent to the design life and expected life of the device, but in no case less than two years from the date of release for commercial distribution by the manufacturer [820.180(b)]. The intent of this requirement is support for potential repairs, corrective actions and recalls. Each manufacturer should make a prudent decision whether to discard records or keep all, or part, of them for a longer period. When requested, distribution records shall be made available to FDA investigators for review and copying during normal business hours. 





DEVICE INSTALLATION





	Section 820.170 on installation requires that each manufacturer establish and maintain adequate installation and inspection instructions and, where appropriate, testing procedures. The purpose of this requirement is to ensure that the device is properly installed and will perform as intended after installation. This regulation applies to medical device systems and complex devices that require set up and adjustment at the location where they are to be used. For example, before a diagnostic x-ray machine can be used, it has to be installed and adjusted and the performance checked. Cardiopulmonary bypass machines also require set up and adjustment at the user location. Manufacturers of such devices shall:





∫	install the device, or have it installed by a representative;


 ∫	inspect and test, as appropriate, the device after installation to assure the device will perform as intended; or


∫	provide adequate instructions and procedures for proper installation by another party.





	These instructions and procedures for proper installation by the manufacturerís representative, user, or third party (820.170) shall include instructions on how to determine that the installed device is safe, performing satisfactorily and ready for use. Safety checks at installation refer to safety aspects directly related to the installation and setup activities and not to intrinsic safety features that have already been checked during final acceptance testing at the factory. 





	The instructions and procedures shall be distributed with the device or otherwise made available to the person installing the device. Such procedures and instructions are part of the device master record and generally include a checklist for the installer to make certain that all necessary installation and checkout activities have been performed correctly. The installer should complete the checklist. If available to the manufacturer, the filled-in checklist or other installation records are part of the device history record.





	Installation and servicing are related activities. Therefore, see Chapter 16, Servicing, for more information.





EXHIBITS





Various forms to show that devices are finished and may be released or stopped from release and a list of some traceable devices are briefly described below and then exhibited.





Finished Product Release Form





This exhibit shows an example of a finished product release form which is actually a checklist for the manufacturing and QC departments of an in vitro diagnostic manufacturer to show that all required processes have been completed. The checklist acts as a reminder of the acceptance forms that are needed for a product and has space for the manufacturing and QC people to indicate that these forms have been completed and reviewed. Finally there is space for the designees to approve or disapprove the lot for release and for comments, if needed.





Release To Finished Goods/Shipping





This exhibit is a release form as described above except that it is for various hardware products. The employee writes in the specification for the product being released.





�
Product Shipping Hold





This exhibit is an example of a form used to stop the shipping of a finished device for reasons related to safety, performance, reliability, regulatory compliance, or other quality requirements.





Release From Product Shipping Hold





This is a form used to release a finished device from a stop shipment order. Because stop orders are always significant, this release form requires a signature by key management.





�Partial List of Traceable Devices





This exhibit lists many of the devices for which a manufacturer must adopt a method of device tracking and the citation to 21 C.F.R. for the device.





It is required that implantable devices and life sustaining devices, whose failure during use as described on the label, could result in significant injury to the user, establish and maintain procedures for identifying with a control number each unit, lot, or batch of finished devices.  Where appropriate, this traceability rule also applies to components. These procedures should facilitate corrective action.  This identification should be documented in the device history record. Many of these devices were formerly called critical devices. 


  �
�






FINISHED PRODUCT RELEASE�



Form No.�



Rev.�



Sheet 1 of 1�
�



Form Approved by:�
�



Date�
�
�



ECN notes:�
�
�
�
�



�
�
�
�
�



Title:  AMYLASE SET�
�
�
�
�



Packaging lot number�



Circle one CATALOG Number  AM-389-01


AM-389-02�
�
�
�



�
�
�
�
�



  The device history documents below were reviewed by Æ


Circle one form number in 2, 5 & 7 below.�
�



  MFG


   ¸�



  QC


   ¸�
�



1. Form # 9926�



Product flow sheet�



�



�
�



2. Form # 1077  or  1078�



Iodine solution�



�



�
�



3. Form # 1082�



Substrate solution�



�



�
�



4. Form # 1083�



Substrate tube filling sheet�



�



�
�



5. Form # 1084  or  1085�



Iodine filling sheet�



�



�
�



6. Form # 1086�



Packaging record�



�



�
�



7. Form # QC-PP-07 or QC-PP-01�



Finished device specification�



�



�
�



�



�



�



�
�



Comments�
�
�
�
�



�
�
�
�
�



�
�
�
�
�



Sign. MFG Designee 	    APP. Yes or No�
�



�



�
�



Comments�
�
�
�
�



�
�
�
�
�



�
�
�
�
�



Signature QC Designee 		 Approved Yes or No�
�
�



�
�






�






RELEASE TO FINISHED GOODS/SHIPPING�



Form # Release-110�



Rev. B�



Sheet 1 of 1�
�



Form Approved by: �
�



Date  4-15-1974�
�
�



ECN notes:�
�
�
�
�



�
�
�
�
�



Type of Product:  [ ] Cable   [ ] Instrument   [ ] Spare/Replacement part/assembly�
�
�
�
�



PRODUCT NAME


�



Part Number�



Serial / Lot Number�
�
�



Inspection Specification Number�
�



Revision�
�
�



PRODUCT STATUS�
�



Circle yes or no Ø�
�
�



1.	Final inspection complete per standards set forth in the QC manual and device inspection specification ?�
�



 YES�



 NO�
�



2.	Device history record packet present ?�
�



 YES�



 NO�
�



If the answer to 1 or 2 is NO, return lot to Production.�
�
�
�
�



3.	Has Final Inspection performed a simulated use test ?�
�



 YES�



 NO�
�



4.	An identified final test data sheet is with the unit ?�
�



 YES�



 NO�
�



�
�



�



�
�



Comments�
�
�
�
�



�
�
�
�
�



�
�
�
�
�



Signature QC Designee	Approved�
�



 YES�



 NO�
�



Quantity Released�
�



 Date�
�
�



�






HOLD NUMBER:�



DATE:�
�



	PRODUCT SHIPPING HOLD


	BY QUALITY ASSURANCE DEPARTMENT�
�
�



The product listed below is on SHIPPING HOLD and under NO circumstances is to be shipped from the factory or any field office without the written approval of the Director, Quality Assurance.�
�
�



PRODUCT:�
�
�



HOLD STARTING DATE:�
�
�



	CLASSIFICATION OF HOLD�
�
�



 [ ] EFFICACY


 [ ] RELIABILITY


 [ ] STERILITY


 [ ] OTHER (describe)�



 [ ] SAFETY


 [ ] REGULATORY COMPLIANCE


 [ ] GOOF UP


�
�



REASON FOR HOLD:�
�
�



�
�
�



�
�
�



�
�
�



ACTION REQUIRED BEFORE RELEASE:�
�
�



�
�
�



�
�
�



�
�
�



MDR/RECALL NOTES:�
�
�



�
�
�



 Sign., DIRECTOR, QUALITY ASSUR�ANCE�
�
�



	DISTRIBUTION:�
�
�



 [ ] Manager, Shipping Depart�ment


 [ ] National Field Manager


 [ ] General Manager


 [ ] Controller


 [ ] V.P. Corporate (if s/e problem)�



 [ ] Director, Manufacturing


 [ ] Division President


 [ ] Director, R & D


 [ ] Director, Marketing


 [ ] V.P. International (if int'l sales)�
�
�






HOLD NUMBER:�



DATE:�
�






	RELEASE FROM PRODUCT SHIPPING HOLD





	BY QUALITY ASSURANCE DEPARTMENT�
�
�



The product listed below is RELEASED from SHIPPING HOLD. 


�
�
�



PRODUCT:


�
�
�



HOLD RELEASE DATE:


�
�
�



REASON FOR RELEASE:


�
�
�






�
�
�






�
�
�






�
�
�






�
�
�



MDR/RECALL NOTES:


�
�
�



�
�
�



�
�
�



 Sign., DIRECTOR, QUALITY ASSUR�ANCE


�
�
�



	DISTRIBUTION:


�
�
�



 [ ] Manager, Shipping Depart�ment


 [ ] National Field Manager


 [ ] General Manager


 [ ] Controller


 [ ] V.P. Corporate (if S/E)�



 [ ] Director, Manufacturing


 [ ] Division President


 [ ] Director, R & D


 [ ] Director, Marketing


 [ ] V.P. International (if Int'l sales)�
�



�
Partial List of Traceable Devices








CFR	Classification


Cite	Name of Device


	





PART 868 -- ANESTHESIOLOGY DEVICES


868.1200	Indwelling blood oxygen partial pressure (PO2) analyzer


868.2375	Breathing frequency monitor


868.5090	Emergency airway needle


868.5160(a)	Gas machine for anesthesia


868.5240	Anesthesia breathing circuit


868.5400	Electroanesthesia apparatus


868.5440	Portable oxygen generator


868.5470	Hyperbaric chamber (Monoplace)


868.5610	Membrane lung for long term pulmonary support


868.5650	Esophageal obturator


868.5720	Bronchial tube


868.5730	Tracheal tube


868.5740	Tracheal/bronchial differential ventilation tube


868.5750	Inflatable tracheal tube cuff


868.5800	Tracheostomy tube and tube off


868.5810	Airway connector


868.5830	Autotransfusion apparatus


868.5895	Continuous ventilator


868.5905	Noncontinuous ventilator (IPPB)


868.5915	Manual emergency ventilator


868.5925	Powered emergency ventilator


868.5935	External negative pressure ventilator





PART 870 -- CARDIOVASCULAR DEVICES


870.1025	Arrhythmia detector and alarm


870.1330	Catheter guide wire


870.1360	Trace microsphere


870.1750	External programmable pacemaker pulse generator


870.1800	Withdrawal-infusion pump


870.3250	Vascular clip


870.3260	Vena cava clip


870.3300	Arterial embolization device


870.3375	Cardiovascular intravascular filter


870.3450	Vascular graft prosthesis of less than 6 millimeters diameter


870.3460	Vascular graft prosthesis of 6 millimeters and greater diameter


870.3470	Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoro-ethylene


870.3535	Intra-aortic balloon and control system


870.3545	Ventricular bypass (assist) device


870.3600	External pacemaker pulse generator


870.3610	Implantable pacemaker pulse generator


870.3620	Pacemaker lead adaptor


870.3650	Pacemaker polymeric mesh bag


870.3670	Pacemaker charger


870.3680	Cardiovascular permanent or temporary pacemaker electrode


870.3700	Pacemaker programmers


870.3710	Pacemaker repair or replacement material


870.3800	Annuloplasty ring


870.3850	Carotid sinus nerve stimulator


870.3925	Replacement heart valve


870.4320	Cardiopulmonary bypass pulsatile flow generator


870.4350	Cardiopulmonary bypass oxygenator


870.4360	Nonroller-type cardiopulmonary bypass blood pump


870.4370	Roller-type cardiopulmonary bypass blood pump


870.5200	External cardiac compressor


870.5225	External counter-pulsating device


870.5300	DC-defibrillator (including paddles)


870.5550	External transcutaneious cardiac pacemaker (noninvasive)


---	Percutaneous transluminal coronary angioplasty (PTCA) balloon dilation catheter


---	Automatic implanted cardioverter defibrillator system





PART 872 -- DENTAL DEVICES


872.3640	Endosseous implant





PART 874 -- EAR, NOSE, AND THROAT DEVICES


872.3620	Ear, nose and throat synthetic polymer material


874.3695	Mandibular implant facial prosthesis


874.3730	Laryngeal prosthesis (Taub design)


874.3820	Emdolmphatic shunt


874.3850	Endolymphaic shunt tube with valve


874.3930	Tympanotomy tube with semipermeable membrane


---	Ear, nose, throat natural polymer-collagen material





PART 876 -- GASTROENTEROLOGY-UROLOGY DEVICES


876.3350	Penile inflatable implant


876.5270	Implanted electrical urinary continence device


876.5540	A-V shunt cannula


876.5630	Peritoneal dialysis system and accessories


876.5820	Hemodialysis system and accessories, dialysate concentrate, hollow fiber capillary dialyzers, disposable dialyzers, high permeability dialyzers, parallel flow dialyzers, single coil dialyzers, twin coil dialyzers, single needle dialysis set, dialysate delivery systems


876.5870	Sorbent hemoperfusion system


876.5880	Isolated kidney perfusion and transport system and accessories


876.5955	Peritoneo-venous shunt


46 FR 7566	Urethral sphincter prosthesis


  (1/23/81)


46 FR 7566	Urethral replacement


  (1/23/81)





PART 878 -- GENERAL AND PLASTIC SURGERY DEVICES





42 FR 63474	Absorbable surgical sutures


  (12/16/77)


42 FR 63474	Nonabsorbable surgical sutures


  (12/16/77)


879.4520	Polytetrafluoroethylene (Teflon) injectable


878.3300	Surgical mesh


878.3500	Polytetrafluoroethylene with carbon fibers composite implant material


878.3530	Inflatable breast prosthesis


878.3540	Silicone gel-filled breast prosthesis


---	Implanted mammary prosthesis of composite saline and gel-filled design


878.3610	Esophageal prosthesis


878.3720	Tracheal prosthesis


878.4300	Implantable clip


878.4750	Implantable staple


---	Maxillofacial prosthesis





PART 880 -- GENERAL HOSPITAL AND PERSONAL USE DEVICES


880.5130	Infant radiant warmer


880.5400	Neonatal incubator


880.5410	Neonatal transport incubator


880.5725	Infusion pump


---	Implanted infusion pump





PART 882 -- NEUROLOGICAL DEVICES


882.5030	Methyl methacrylate for aneurysmorrhaphy


882.5150	Intravascular occluding catheter


882.5200	Aneurysm clip


882.5225	Implanted malleable clip


882.5250	Burr hole cover


882.5300	Methyl methacrylate for cranioplasty


882.5320	Preformed alterable cranioplasty plate


882.5330	Preformed nonalterable cranioplasty plate


882.5360	Cranioplasty plate fastener


882.5550	Central nervous system fluid shunt and components


882.5820	Implanted cerebellar stimulator


882.5830	Implanted diaphragmatic/phrenic nerve stimulator


882.5840	Implanted intracerebral/subcortical stimulator for pain relief


882.5850	Implanted spinal cord stimulator for bladder evacuation


882.5860	Implanted neuromuscular stimulator


882.5870	Implanted peripheral nerve stimulator for pain relief


882.5880	Implanted spinal cord stimulator for pain relief


882.5880	Epidural spinal electrode


882.5900	Preformed craniosynostosis strip


882.5910	Dura substitute


882.5950	Artificial embolization device


---	Lyophilized human (cadaver) dura mater


---	Stabilized epidural spinal electrode


---	Implanted intracranial pressure monitor


---	Totally implanted spinal cord stimulator for pain relief





PART 884 -- OBSTETRICAL AND GYNECOLOGICAL DEVICES


884.5360	Contraceptive intrauterine device (IUD) and introducer


884.5380	Contraceptive tubal occlusion device (TOD) and introducer





PART 886 -- OPHTHALMIC DEVICES


886.3300	Absorbable implant (scleral buckling method)


886.3400	Keratoprosthesis


886.3600	Intraocular lens


886.3920	Eye valve implant





PART 888 -- ORTHOPEDIC DEVICES


888.3000	Bone Cap


888.3010	Bone fixation cerclage


888.3020	Intramedullary fixation rod


888.3025	Passive tendon prosthesis


888.3027	Polymethyllmethaccrylate (PMMA) bone cement


888.3030	Single/multiple component metallic bone fixation appliance and accessories


888.3040	Smooth or threaded metallic bone fixation fastener


888.3050	Spinal interlaminal fixation orthosis


888.3060	Spinal intervertebral body fixation orthosis


888.3100	Ankle joint metal/composite semi-constrained cemented prosthesis


888.3110	Ankle joint metal/polymer semiconstrained cemented prosthesis


888.3120	Ankle joint metal/polymer non-constrained cemented prosthesis


888.3150	Elbow joint metal/metal or metal/polymer constrained cemented prosthesis


888.3160	Elbow joint metal/polymer semi-constrained cemented prosthesis


888.3170	Elbow joint radial (hemi-elbow) polymer prosthesis


888.3180	Elbow joint humeral (hemi-elbow) metallic uncemented prosthesis


888.3200	Finger joint metal/metal constrained uncemented prothesis


888.3210	Finger joint metal/metal constrained cemented prosthesis


888.3220	Finger joint metal/polymer constrained cemented prosthesis


888.3230	Finger joint polymer constrained prosthesis


888.3300	Hip joint metal constrained cemented or uncemented prosthesis


888.3310	Hip joint metal/polymer constrained cemented or uncemented prosthesis


888.3320	Hip joint metal/metal semi-constrained, with a cemented acetabular component, prosthesis


888.3330	Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis


888.3340	Hip joint metal/composite semi-constrained cemented prosthesis


888.3350	Hip joint metal/polymer semi-constrained cemented prosthesis


888.3360	Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis


888.3370	Hip joint (hemi-hip) acetabular metal cemented prosthesis


888.3380	Hip joint femoral (hemi-hip) trunnion-bearing metal/polyacetal cemented prosthesis


888.3390	Hip joint femoral (hemi-hip metal/polymer cemented or uncemented prosthesis


888.3400	Hip joint femoral (hemi-hip) metallic resurfacing prosthesis


888.3410	Hip joint metal/polymer semi-constrained resurfacing cemented prosthesis


888.3480	Knee joint femorotibial metallic constrained cemented prosthesis


888.3490	Knee joint femorotibial metal/composite non-constrained cemented prosthesis


888.3500	Knee joint femorotibial metal/composite semi-constrained cemented prosthesis


888.3510	Knee joint femorotibial metal/polymer constrained cemented prosthe�sis


888.3520	Knee joint femorotibial metal/polymer non-constrained cemented prosthesis


888.3530	Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis


888.3540	Knee joint patellofemoral polymer/metal semi-constrained cemented prosthesis


888.3550	Knee joint patellofemorotibial polymer/metal/metal constrained cemented prosthesis


888.3560	Knee joint patellofemorotibial polymer/metal/polymer semi-con�strained cemented prosthesis


888.3570	Knee joint femoral (hemi-knee) metallic uncemented prosthesis


888.3580	Knee joint patellar (hemi-knee) metallic resurfacing uncemented prosthesis


888.3590	Knee joint tibial (hemi-knee) metallic resurfacing uncemented prosthesis


888.3640	Shoulder joint metal/metal or metal/polymer constrained cemented prosthesis


888.3650	Shoulder joint metal/polymer non-constrained cemented prosthesis


888.3660	Shoulder joint metal/polymer semi-constrained cemented prosthesis


888.3680	Shoulder joint glenoid (hemi-shoulder) metallic cemented prosthesis


888.3690	Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis


888.3720	Toe joint polymer constrained prosthesis


888.3730	Toe joint phalangeal (hemi-toe) polymer prosthesis


888.3750	Wrist joint carpal lunate polymer prosthesis


888.3760	Wrist joint carpal scaphoid polymer prosthesis


888.3770	Wrist joint carpal trapezium polymer prosthesis


888.3780	Wrist joint polymer constrained prosthesis


888.3790	Wrist joint metal constrained cemented prosthesis


888.3800	Wrist joint metal/polymer semi-constrained cemented prosthesis


888.3810	Wrist joint ulna (hemi-wrist) polymer prosthesis


Section Fourteen; Page � PAGE �1� of � NUMPAGES �13�














